First test.
        The different samples of the O-ring lie at the copper strip sample and clamped between two G-10 pieces with clamps. Three types of clamps provided different pressure high,medium and low. Clamps with samples situated inside of the vacuum chamber. The hole was drilled in the dome of vacuum chamber and the gas output tube was glued with the green epoxy. The lead brick lied on the top of the chamber and provided the pressure at the EPDM spacer between dome and bottom of the chamber. Gas was connected to input tube at the bottom of the vacuum chamber. Gas mixture was Ar/CO2 40%/60% respectively. The cup of water was inside of the gas volume and provided humidity by evaporating. Vacuum chamber was heating with two lamps up to 90C degrees. The temperature determined by thermometer which also be inside of the gas volume.
The first test started at August 8 2012 and run for five days. Vacuum chamber was heating by the lamp three days during working time (8a.m.-6p.m.). Five clamps with samples was inside of the chamber during test. It was EPDM with high pressure clamp, EPDM with medium pressure clamp, EPDM with low pressure clamp, Viton with high pressure clamp, Viton plus Apiezon with high pressure clamp. Also five addition copper strip samples was lie at the EPDM spacer. Copper side was in contact with EPDM and divided at two parts by vacuum chamber dome edge. So one part of the samples was in the air another part inside gas volume. Pic.1.

        After chamber was opened EPDM samples clearly shown the corrosive damage, also at the Viton sample o-ring trace was visible. Viton plus Apiezon did not show corrosive effect. All samples was moved to SEM analysis.

        Result of SEM analysis

        EPDM sample was divide into two little samples, also one sample was use as reference.
   Analysis was taken into five different areas. Spectrum analysis of elements shown at Pic. All present spectrum and photos was made with 1K magnification.

   The biggest concentration of sulfur was found at the areas where edges of o-ring was(area2,area4). Pictures from SEM also clearly shows big concentration of sulfur crystals at the edges areas.
        Viton sample was also analyzed.

        Second test.
Second test was started at August 14 and include 5 samples. Viton+Grease, EPDM + Gr,
Viton, EPDM (high pressure), EPDM(low pressure). All this samples was heated by the lamp up to 90C degrees for 7 days. No water inside. Same corrosive picture after opening. Samples was moved to SEM analysis.

        Third test.

        Third test was start at August 21. It was one Long strip + Viton/Grease with high pressure clamp.
Water was inside and evaporate completely. Also Test chamber stay for two day after it and all weekend. Chamber was opened at August 29. Result shown same corrosive traces.

     Fourth test.

      To simulate real pressure between O-ring material and copper strips gas spacer was use instead of g10 pieces. O-ring material with apiezon lie into the grove like in real situation. Only viton sample was test. Clamps was too big and lead break provide pressure instead.

The dome of vacuum chamber was cut in two places. It was made for gas spacer. The gas spacer lied at edges at bottom of vacuum chamber. Addition gap between gas spacer and rubber was excluded by adding one more rubber cell. Putty used for holes to avoid leaks. During this test lamp was turn on during night too. The lead break inside absorb heat a lot, because of it temperature inside was 50 degree of Celsius. Also one lamp was burn during test. Water inside chamber evaporated at 4/5 of original volume. Chamber was assembled in Friday and test was run during weekend. On Monday one lamp was burned. The result was the same as in test before. Visible traces from O-ring was found.

        Fifth test.

The clamps was cut, after it was enough place for them inside of gas volume. The test with Viton+Grease o-ring and copper strips which lie perpendicular to o-ring was assembled. Assemble with clamps helps to increase temperature inside. September 2 to 6 ???

        Sixth test.

    Next test was with copper samples cover with Hysol epoxy width approximately 6mkm . The same construction was closed for 5 days. Temperature inside was 57 C but thermometer was assembled no good and indicate temperature wrong. 16 ml of water was evaporate completely at second day.

    The result shown no different between copper and copper covered with Hysol samples. Both of them had corrosive traces.