Hall D Wire Characterization

Samples F7-2 & F7-1

- Sample preparation: ultrasonic cleaning in DI water for 45 minutes.

- Sample installation: for SEM inspection wires were installed on a sample holder (diameter=36 mm) on 2 stripes of conductive carbon tape, 2 fragment of each sample per stripe (nomenclature: F5-1a, F5-1b and F5-2a, F5-2b).

Pic1. Sample installation.

Wire inspection:

- Each wire fragment was examined along its length (beginning at approximately 5mm from the sample holder edge and ending ~5mm before the opposite edge) at magnification of ~x600, at fast scan rate.

- Images at designated magnifications (x300, x500, x800, x1000) were taken randomly along the sample length;

- Images for the ovality measurements were taken at 1000x magnification in three points of each wire fragment: in the middle (approximately) and close to both ends;

- Measurements of wire diameter, as visible on the images, were done using the Quartz PCI Image Management System, 3 measurements per image;

- Additional images at various magnifications were taken at the points of interest (variations in sample topography, contaminated areas, etc.)

- EDS analysis was performed on the most typical points of interest.

Note:

Sample F7-1 was inspected BEFORE the current protocol was established, thus the report on F7-1 differs slightly from other reports.

Sample F7-2 A, B

I. Images at designated magnifications

| F7-2A-1 | F7-2B-1 |
| --- | --- |
| Magnification: 300x | Magnification: 300x |
| F7-2A-2 | F7-2B-2 |
| Magnification: 500x | Magnification: 500x |

| F7-2A-3 | F7-2B-3 |
| --- | --- |
| Magnification: 800x | Magnification: 800x |
| F7-2A-4 | F7-2B-4 |
| Magnification: 1000x | Magnification: 1000x |

II. Ovality Measurements

| F7-2A-1o | F7-2B-1o |
| --- | --- |
| Magnification: 1000x | Magnification: 1000x |
| F7-2A-2o | F7-2B-2o |
| Magnification: 1000x | Magnification: 1000x |
| F7-2A-3o | F7-2B-3o |
| Magnification: 1000x | Magnification: 1000x |

III. Points of Interest

| F7-2A-3 | F7-2A-3 |
| --- | --- |
| Magnification: 6000x | EDS Analysis. Environmental contamination |
| F7-2A-5 Minor surface damage | F7-2A-6 Minor surface damage |
| Magnification: 2490x | Magnification: 1800x |
| F8-2A -8 | F7-2A -8 |
| Magnification: 30000x | EDS Analysis. Surface damage. Possible traces of Cu |

| F7-2A-7 | F7-2A-7 |
| --- | --- |
| Magnification: 50000x | EDS Analysis. Surface damage. Possible Cl contamination |
| F7-2B-5 | F7-2B-8 |
| Magnification: 2500x Contamination | Magnification: 30000x |

Sample F7-1 A, B

Images at designated magnifications

| F7-1A-1 | F7-1B-1 |
| --- | --- |
| Magnification: 250x | Magnification: 250x |
| F7-1A-2 | F7-1B-2 |
| Magnification: 500x | Magnification: 500x |

| F7-1A-3 | F7-1B-3 |
| --- | --- |
| Magnification: 800x | Magnification: 800x |
| F7-1A-4 | F7-1B-4 |
| Magnification: 1000x | Magnification: 1000x |

II. Ovality Measurements

| F7-1A -1o | F7-1B -1o |
| --- | --- |
| Magnification: 1000x | Magnification: 1000x |
| F7-1A -2o | F5-2B-2o |
| Magnification: 1000x | Magnification: 1000x |
| F7-1A -3o | F7-1B -3o |
|  | Sorry, image was not saved. |
| Magnification: 1000x | Magnification: 1000x |

III. Points of Interest

| F7-1A-7 Surface contamination. | F7-1A-7 |
| --- | --- |
| Magnification: 1000x | EDS Analysis. Surface contamination. Carbon. |
| F7-1A-4 Surface structure detail. | F7-1A-minor surface scratch |
| Magnification: 100000x | Magnification: 1500x |
| F7-1B-8 Surface feature. | F7-1B-11 Surface feature. |
| Magnification: 15000x | Magnification: 100000x |

Note: No Cu was observed by EDS probe for the features F7-1B-8 and F7-1B-11.