## Wireboards
| Wireboard # | Epoxy | Humi-seal | Replaced Capacitors | Wire Deadened | Wireboard # | Epoxy | Humi-seal | Replaced Capacitors | Wire Deadened | Dark Current (nA) | Stability
Spare | W001 | Epo-lite | No | No | No | W001 | Epo-lite | No | No | No | >1000 @ 1900v | No
Spare | W002 | Epo-lite | Yes | No | Yes | W002 | Epo-lite | Yes | No | 6cm | 130 | Yes
Spare | W003 | Epo-lite | No | No | No | W003 | Epo-lite | No | Removed | No | 750 | Yes
Spare | W006 | Epo-lite | Yes | No | No | W006 | Epo-lite | Removed | Removed | No | 1200 | Yes
Package 1 | W005 | Epon | Yes | No | Yes | W0027 | Epon | No | Yes | 7.8cm | 1200 | Yes
| W008 | Epon | No | Yes | Yes | W0028 | Epon | No | Yes | 7.8cm | 350 | Yes
| W009 | Epon | No | Yes | Yes
| W010 | Epon | No | Yes | Yes
| W011 | Epon | No | Yes | Yes
| W012 | Epon | No | Yes | Yes
Package 2 | W007 | Epon | No | Yes | Yes
| W014 | Epon | No | Yes | Yes
| W015 | Epon | No | Yes | Yes
| W017 | Epon | No | Yes | Yes
| W019 | Epon | No | Yes | Yes
| W016 | Epon | No | Yes | Yes
Package 3 | W020 | Epon | No | Yes | Yes
| W021 | Epon | No | Yes | Yes
| W022 | Epon | No | Yes | Yes
| W023 | Epon | No | Yes | Yes
| W004 | Epon | No | Yes | Yes
| W013 | Epon | No | Yes | Yes
Package 4 | W024 | Epon | No | Yes | Yes
| W025 | Epon | No | Yes | Yes
| W026 | Epon | No | Yes | Yes
| W018 | Epon | No | Yes | Yes
| W027 | Epon | No | Yes | Yes
| W028 | Epon | No | Yes | Yes

## Current Tracking
FDC Construction Tracking | [8 empty] | 2012-08-16
FDC1 | UP | End Window | EW003
| Cell C6 | Type 3 Cathode | C016 | Wireboard | W005 | Gas Spacer | Type 1 Cathode | C018
| Cell C5 | Type 2 Cathode | C019 | Wireboard | W009 | Gas Spacer | Type 1 Cathode | C020
| Cell C4 | Type 2 Cathode | C021 | Wireboard | W008 | Gas Spacer | Type 1 Cathode | C022
| Cell C3 | Type 2 Cathode | C023 | Wireboard | W011 | Gas Spacer | Type 1 Cathode | C026
| Cell C2 | Type 2 Cathode |  | Wireboard | W010 | Gas Spacer | Type 1 Cathode
| Cell C1 | Type 2 Cathode |  | Wireboard | W012 | Gas Spacer | Type 3 Cathode | C017
| Down | End Window | EW004
FDC2 | UP | End Window | EW005 | [6 empty] | 90%
| Cell C6 | Type 3 Cathode | C061 Daughter Cards (C015) | Wireboard | W007 | Gas Spacer | Type 1 Cathode | C060 Daughter Cards (New Frame)
| Cell C5 | Type 2 Cathode | C062 Daughter Cards (C013) | Wireboard | W014 | Gas Spacer | Type 1 Cathode | C063 Daughter Cards (C014)
| Cell C4 | Type 2 Cathode | C064 Copper Flpas (C009) | Wireboard | W015 | Gas Spacer | Type 1 Cathode | C066 Glue Foils
| Cell C3 | Type 2 Cathode | C067 Glue Foils | Wireboard | W017 | Gas Spacer | Type 1 Cathode
| Cell C2 | Type 2 Cathode |  | Wireboard | W019 | Gas Spacer | Type 1 Cathode
| Cell C1 | Type 2 Cathode |  | Wireboard | W016 | Gas Spacer | Type 3 Cathode | C065 Glue Frame (C003)
| Down | End Window | EW006
FDC3 | UP | End Window | EW007 | [6 empty] | 100%
| Cell C6 | Type 3 Cathode | C030 | Wireboard | W020 | Gas Spacer | Type 1 Cathode | C027
| Cell C5 | Type 2 Cathode | C036 | Wireboard | W021 | Gas Spacer | Type 1 Cathode | C032
| Cell C4 | Type 2 Cathode | C038 | Wireboard | W022 | Gas Spacer | Type 1 Cathode | C035
| Cell C3 | Type 2 Cathode | C031 | Wireboard | W023 | Gas Spacer | Type 1 Cathode | C037
| Cell C2 | Type 2 Cathode | C034 | Wireboard | W004 | Gas Spacer | Type 1 Cathode | C040
| Cell C1 | Type 2 Cathode | C041 | Wireboard | W013 | Gas Spacer | Type 3 Cathode | C029
| Down | End Window | EW008
FDC4 | UP | End Window | EW009 | [6 empty] | 88%
| Cell C6 | Type 3 Cathode | W043 | Wireboard | W024 | Gas Spacer | Type 1 Cathode | C044
| Cell C5 | Type 2 Cathode | C046 | Wireboard | W025 | Gas Spacer | Type 1 Cathode | C048
| Cell C4 | Type 2 Cathode | C045 | Wireboard | W026 | Gas Spacer | Type 1 Cathode | C049
| Cell C3 | Type 2 Cathode | C055 | Wireboard | W018 | Gas Spacer | Type 1 Cathode | C050
| Cell C2 | Type 2 Cathode | C051 | Wireboard | W028 Testing | Gas Spacer | Type 1 Cathode | C052
| Cell C1 | Type 2 Cathode | C057 | Wireboard | W027 Testing | Gas Spacer | Type 3 Cathode | C042
| Down | End Window | EW010
Spare |  | End Window | EW001 | [6 empty] | 92%
|  | Type 3 Cathode |  | Wireboard | W002 | Gas Spacer | Type 3 Cathode | C054
|  | Type 2 Cathode | C047 | Wireboard | W001 | Gas Spacer | Type 1 Cathode | C056
|  | Type 2 Cathode | C053 | Wireboard | W003 | Gas Spacer | Type 1 Cathode | C058
|  | Type 2 Cathode | C059 | Wireboard | W006 | Gas Spacer | Type 1 Cathode
|  | End Window | EW002
Nonconforming | FDC Componant |  | Nonconformance |  | Corrective Action
| W001 |  | Trips HV power supply |  | This board was cleaned in the ultrasonic cleaner. After cleaning the board would not hold HV. The board was placed in a vacuum chamber to remove moisture. It is theorized that the flux that was removed during the ultrasonic cleaning formed a film and was deposited on the board as it was removed.
| W002 |  | Sector 4 trips HV power supply |  | There is a high resistance path (~30-40 mohm) between -HV4 and +HV4. F86 and S86 were replaced due to contamination from the copper sulfate acid mixture used in the deadening process that has dripped on this spot. After thoroughly cleaning the area and replacing the wires the path was still there. The board is being conditioned and is making progress. Might be usable as a spare.
| C047 |  | Cut trace on channel #1 |  | During a repear to this cathode a damaged trace was coated with a small amount of Hysol. The Hysol did not allow the wireboard to lay flat on the cathode. While attempting to remove some of the Hysol the trace was damaged. We will try to repair with conductive TIGA920 conductive epoxy. Reparied the cathode and got a reading of <4ohms. Board returned to service.
| W013 |  | Too much heat applied to pads during capacitor removal |  | Specified soldering tip to be used during capacitor installation and removal. Board is in the EEFAB Department to be fixed. Capacitors were resoldered to the pads and then epoxied down using Scotchweld 1838. Board returned to service.
| W004 |  | S17 was broken at the pad (possibly during cleaning or testing) |  | Replace S17 and retest. Board returned to service
| C033 |  | Type 2 Foil Misaligned on G10 Frame |  | Retrained techs on procedure and reiterated the importance of high quality
| C039 |  | Misaligned on Tension Ring |  | Retrained techs on procedure and reiterated the importance of high quality
| W018 |  | G10 Indexed one hole off |  | Vision Machine is reworking G10 frame by adding counter bore to bolt holes. Board returned to service.
| C003 |  | Copper Corrosion in o-ring area
| C008 |  | Copper Corrosion in o-ring area |  | Reworked frame for use in C064
| C010 |  | Copper Corrosion in o-ring area
| C011 |  | Copper Corrosion in o-ring area
| C013 |  | Copper Corrosion in o-ring area |  | Reworked frame for use in C062
| C015 |  | Copper Corrosion in o-ring area |  | Reworked frame for use in C061
| C006 |  | Copper Corrosion in o-ring area
| C007 |  | Copper Corrosion in o-ring area
| C009 |  | Copper Corrosion in o-ring area
| C012 |  | Copper Corrosion in o-ring area
| C014 |  | Copper Corrosion in o-ring area |  | Reworked frame for use in C063
| C004 |  | Copper Corrosion in o-ring area
| W023 |  | S96 has a short to ground between C96 and the via ajacent to R96 (0402). |  | Drilled out the via and determined the short is between the cap and the via. We will have to mill out the G10 and Rohacell and follow the trace until the short is found, then a jumper wire was installed. Board returned to service.
Complete
In-Process
Copper traces need repairing
Nonconforming
[7 empty] | X | Assembled In Package
Note on G3: debutler: Stacked 8/31/2011
Note on G4: debutler: Stacked 9/15/2011
Note on G5: debutler: stacked 9/27/2011
Note on G6: debutler: Stacked 9/27/2011
Note on G7: debutler: Stacked 9/30/2011
Note on G8: debutler: Stacked 10/12/2011
Note on G19: debutler: Stacked on 4/18/2012
Note on G20: debutler: Stacked 4/25/2012
Note on G21: debutler: Stacked 4-27-2012
Note on G22: debutler: Stacked 4-27-2012
Note on G23: debutler: Stacked 5-10-2012
Note on G24: debutler: Stacked 5-10-2012
Note on G27: debutler: Stacked 5-31-2012
Note on G28: debutler: Stacked 5-31-2012
Note on G29: debutler: Stacked 6-1-2012
Note on G30: debutler: Stacked 6/20/2012
Note on G31: debutler: Stacked 6/22/2012
Note on G32: debutler: Stacked 6/25/2012

## Package Repair
FDC Repair Construction Tracking | [11 empty] | 2012-08-16
FDC1 | Stage I | [11 empty] | 0%
| UP | End Window | EW003 | 1
| Cell C6 | Type 3 Cathode | C003 |  | Wireboard | W005 | 1 | Gas Spacer | 2 | Type 1 Cathode | C006 | 1
| Cell C5 | Type 2 Cathode | C008 |  | Wireboard | W009 | 1 | Gas Spacer | 2 | Type 1 Cathode | C007 | 1
| Cell C4 | Type 2 Cathode | C010 |  | Wireboard | W008 | 1 | Gas Spacer | 2
|  | Type 3 Cathode | C004
|  | End Window | EW004 | 1
Stage II
| Cell C3 | Type 2 Cathode | C011 |  | Wireboard | W011 | 1 | Gas Spacer | 2 | Type 1 Cathode | C009 | 1
| Cell C2 | Type 2 Cathode | C013 |  | Wireboard | W010 | 1 | Gas Spacer | 2 | Type 1 Cathode | C012 | 1
| Cell C1 | Type 2 Cathode | C015 |  | Wireboard | W012 | 1 | Gas Spacer | 2 | Type 1 Cathode | C014 | 1
Down
Notes: Package has been moved inside the cleanroom and all pre-amps and cooling loops have been removed. The strong back has also been removed and the package is on the granite table.
FDC2 | Stage I | [11 empty] | 0%
| UP | End Window | EW005 | 1
| Cell C6 | Type 3 Cathode | C016 |  | Wireboard | W007 | 1 | Gas Spacer | 2 | Type 1 Cathode | C027 | 1
| Cell C5 | Type 2 Cathode | C019 |  | Wireboard | W014 | 1 | Gas Spacer | 2 | Type 1 Cathode | C032 | 1
| Cell C4 | Type 2 Cathode | C021 |  | Wireboard | W015 | 1 | Gas Spacer | 2
|  | Type 3 Cathode | C017
|  | End Window | EW006 | 1
Stage II
| Cell C3 | Type 2 Cathode | C011 |  | Wireboard | W017 | 1 | Gas Spacer | 2 | Type 1 Cathode | C035 | 1
| Cell C2 | Type 2 Cathode | C013 |  | Wireboard | W019 | 1 | Gas Spacer | 2 | Type 1 Cathode | C037 | 1
| Cell C1 | Type 2 Cathode | C015 |  | Wireboard | W016 | 1 | Gas Spacer | 2 | Type 1 Cathode | C040 | 1
Down
Notes: Package is in the storarage rack and will be refirbished after package 1.
Nonconforming | FDC Componant |  | Nonconformance |  |  | Corrective Action
Complete
In-Process
Nonconforming