Solenoid coil instrumentation summary

Aug 18, 2005

Coil Electrical Status  8-16-2005

Coil electrical measurements

| Coil # | Across Coil | LH Lead - Ground | RH Lead - Ground |
| --- | --- | --- | --- |
| 1 | 4.9 O | 2.2 O | 6.4 O |
| 2 | 3.2 O | OPEN | OPEN |
| 3 | 2.7 O | 2.6 O | 0.2 O |
| 4 | 5.2 O | OPEN | OPEN |

Coil Inductances

Coil 1        Coil 2        Coil 3     Coil 4

372 mH    244 mH    172 mH   763 mH

Note:  Coil 3 short repair in progress.

Internal instrumentation and voltage taps      Status     8-16-2005

Each coil has voltage taps(VT) and Carbon Resistance Thermometers(CRT) in the Helium vessel and Thermalcouples(TC) on the N2 shield and strain gauges(SG) on the support posts.

Coil 1            Coil 2          Coil 3         Coil 4

Voltage Taps                   OK                  OK             OK              OK

CRT (LHE)                7 of 8 OK             OK             OK         4 of 7 OK

TC  (LN2)               Removed            Removed      TBD        Removed

TC (Helium)            Some TC’s remain inside the coil cases(LANL report)

PT- 102 (LN2)           16 New             16 New          TBD        30 New

SG (new)                      6 New               6 New          TBD         6 New

SID                           Some Silicon Diodes are in coil cases see LANL report

Notes

CRT --- Carbon Resistance Thermometer 300 Ohms    4 K to 300 K

Location as in LANL report

TC    --- Thermalcouple(see LANL report)                   80 K to 300 K

PT    --- Platinum Resistance Thermometer                  40 K to 300 K

Locations for each coil ( one primary plus one redundant at each)

LN2 fill    (6 o’clock)

LN2 vent   (12 o’clock)

#1 Outer shield    (3 o’clock)

#2 Outer shield    (9 o’clock)

#1 Inner shield

#2 Inner shield

#1 Intermediate shield

#2 Intermediate shield

SG    --- Strain Gauge-located 4 on axial supports and two on radial supports

VT-     All are OK, locations as in LANL report and “short report”.