# MEMORANDUM

To:	Distribution

From:	F. Dylla

Subject:	FEL Upgrade Project Weekly Brief – February 2-6, 2004

Date:	February 6, 2004

Highlights:

We had a very successful week with commissioning of the 10 kW upgrade. The goal for the week was to achieve lasing at 6 microns with the high Q optical cavity configuration (i.e., using the 2% outcoupling optics), so that we could measure the laser gain at 6 microns and measure the absorption loss in this outcoupler. We were successful on all accounts and then some:

- We successfully lased at 6 microns producing > 1.5 kW cw of outcoupled light with the driver accelerator operating at 88.5 MeV and 3 mA.

- During this lasing the intracavity laser power was 56 kW which is more than half the intracavity power that will be circulating when 10 kW will be extracted from the low Q (10 % outcoupler) cavity configuration.

- While the 56 kW of intracavity is circulating, we see no evidence of a drop in efficiency which would be indicative of excessive loading on the optics; in addition half of the mirror figure control current on the high reflector (HR) was needed to handle this power, indicating we should have the control we need for operation at or near 10 kW.

- With the available high power 6 micron light, we were able to measure the total optical loss in the 2% outcoupling optic actually, (2.7-2.8%); a value of  ~550ppm was obtaining which translates into the ability to couple out approximately 8.5 kW when we install the 10% outcoupling optic (assuming the same coating loss.)

(Note:  we had asserted at the Semi-Annual Review that we would have lower loss optics at 6 microns than our previous attempts with 10 micron optics which showed absorption greater than 2500ppm. We have additional 6 micron optics on order from multiple vendors).

- We were able to boost the linac energy 10% to 88.5 MeV.

- By accounting for the newly discovered gradient imbalance in the injector cryounit we achieved close agreement between our injector performance and injector modeling with the PARMELA code.

- We obtained lasing gain  measurements at 6 microns to compare with our codes and we made substantial progress in understanding the non-linear effects in the electron beam transport.

- The third cryomodule for the FEL linac is ready for transport and placement in the FEL vault on Monday morning.

Details on these and other accomplishments for the week are given below.

NOTE: HOLD THESE DATES:

The 2004 FEL Users/Laser Processing Consortium Workshop will be held at Jefferson Lab on March 10-11, 2004.

Commissioning Summary:

This past week has been a very good one.  We learned a great deal about the non-linear transport of the accelerator and started to get much better agreement between PARMELA and the injector performance.  We also were able to push power up to half of the circulating level required for 10 kW operation without seeing any drop in efficiency.

During the last couple of days of last week we worked on "window washing" in the cryounit and were able to push the power delivered to the waveguide to over 20 kW.  Note that the effective power for this setup is quite high since most of the power is reflected back into the circulator and creates standing wave anti-nodes with very high fields.

On Monday we started on the push to achieve 10 kW at 6 microns.  To raise the gain as high as possible we opted to raise the linac energy as high as possible.  By operating almost all cavities at their "drive high" values, we were able to reach 88.5 MeV when operating 10 degrees off crest.  We did a careful match of the beam through the linac to the wiggler at this energy.  We also changed the backleg setup to have a full 1.5 periods of betatron phase instead of the single period we have been operating at recently.  This seems to have shorted the electron bunch.  We believe that this is due to cancellation of non-linear terms in the first arc.

We also did an experiment to try to measure the relative gradients in the two cavities in the cryounit.  Results of this experiment indicate that the first cavity is 10% lower gradient than EPICs is saying.  To get the injector energy this implies that the second cavity is running 10% higher than EPICs is saying.  When these values are put into PARMELA the behavior is strikingly close to what we see in the experiment, including the fact that the beam is converging at the exit of the cryounit.  We will use this information to re-optimize the injector over the next week.  We may want to balance the gradients in order to reduce the power in the second SRF cavity.

On Tuesday we measured the dispersion, both linear and non-linear, in the backleg.  The non-linear dispersion may be the source of emittance growth between the linac and the backleg.  We found last week that the emittance growth was the same at full charge and half charge so it is not a collective effect.  Using the new electron beam setup at 88.5 MeV we were able to lase at 6 microns on Tuesday afternoon with the 98% output coupler.

On Wednesday we optimized the laser at 6 microns and got a detuning curve as long as 16 microns.  This is close to what one expects from the model.  When we measured the gain it was only 21% compared to a prediction of 30%.  We did find that the gain was quite sensitive to the setup but the power was not so it may not have been optimized when we measured it.   The resonator losses were almost exactly what we expected, indicating that the losses were not extremely high.  We also improved the match to the linac so that we could run up the power with 6% duty cycle.  We found that we could not run at the peak of the detuning curve without loss due to transverse beam blowup in the linac.  This is due to the use of a single family of sextupoles in the second arc.  The energy spread was at least 10% in arc 2 and this high an energy spread requires two families of sextupoles.

Sextupole correction of aberrations provides improved beam quality at the wiggler and control of beam loss during energy recovery. A single-family sextupole pair is required for minimization of bunch length at the FEL; a second single-family is needed to adjust bunch length prior to energy recovery so as to provide sufficient energy compression to remain within the machine energy aperture.

Use of a single family does not, however, allow correction of variations of beam position with momentum. As a consequence, beam quality at the wiggler is compromised by uncompensated second order dispersion, with spot size growth on the order of 10% from this source [1]. This corresponds to an emittance dilution of order 20% – tolerable, though nonideal.

Given recent progress in FEL high power performance, this effect imposes however a more severe limitation during energy recovery. While lasing strongly, an energy spread of order ±5% is imposed on the beam. Inherent lattice aberrations couple to the large momentum spread, leading to both betatron mismatch and shifts in the central orbit position while traversing the second arc. The former can be managed by an appropriate selection of the betatron matching solution, but the latter is manageable solely through sextupole correction. Figure 1 presents beam spots at three locations in the machine for various strengths of lasing. Growth in spot size is linear at the first location – in the center of the 180o dipole – but is not linear at the latter two places, immediately following the energy recovery arc and midway through the reinjection telescope. This spot size growth is consistent with both analytic estimates and DIMAD simulation of machine operational practices and set points.

On Thursday, using the new frame grabber system, we characterized the non-linear dispersion beam growth in the second arc and the extraction efficiency of the laser.  The efficiency was not as high as expected (0.7-0.8%), probably due to a non-optimized laser setup.  We then set up for CW lasing and slowly ramped the current up until 1.5 kW was achieved (see figure 2).  Note that this power output implies 56 kW of circulating power in the optical cavity.  The mirrors seem to have survived this with no damage and the efficiency, though low (500-600 W/mA), was independent of power up to 1.5 kW.  This is a good indication that power levels near 10 kW are achievable with a 90% output coupler if the laser efficiency and gain can be optimized.  Losses in the mirrors, measured by the cooling water temperature rises, show losses consistent with operation at over 8 kW.  We would have raised the power up higher but we had a vacuum incident when the Brewster window through which we were sending the beam cracked, letting the optical cavity and the second arc to high enough pressure that we could not restart the ion pumps.  Since we were scheduled to shut down anyway we stopped there instead of fixing the window (see figure 3).

The next week will be dedicated to getting the new cryomodule into its commissioning position and getting the accelerator put back together.  We should then have a 90% reflective output coupler for 6 micron operation available and can try to push the power up to over 5 kW.

Figure  1: “Wes-cam” [2] screenshots of beam at center of 2nd 180o dipole, exit of energy recovery arc, and midway through reinjection telescope, upstream of beam merge.

## Notes

[1] Dispersion measurements with ±2% offset present a systematic orbit oscillation of order 1 mm independent of the sign of the energy change at locations with spot sizes of order 3  mm. The orbit oscillation decays more rapidly than linearly with decreases in energy. At the nominal beam energy spread of  ±1%, this corresponds to ~1/4 mm, or 10%, change in spot size

[2] Courtesy of Wes Moore, with assistance from Steve Benson and a reminder from Kevin Jordan about the viewer scales.

Figure 2. CW operation at >1.5 kW using high Q resonator.  This corresponds to 56 kW of circulating power in the resonator.

Figure 3. Exit Brewster window after vacuum incident.  The 1.5 kW beam may not have been centered in the window, leading to fracture of the window.  It has withstood over 2 kW at 3 µm in the IR Demo.

Management:

We developed a detailed plan for reaching 10 kW with the Upgrade based on the preliminary plan we presented at last week’s project semiannual review and the feedback from the review panel.  The plan involves interlinking three efforts: (1) the present mainline effort involving the testing of lower loss optics with the present 2 cryomodule configuration; and the commissioning and use of the 3rd cryomodule to boost the linac energy by 2x; (2) two options for modified wigglers that could be incorporated within 12 weeks; and (3) and several low cost but high technical risk options for enhancing 10 kW output.   A technical analysis of these options including cost/schedule implications was completed this week and is being forwarded to the DOE and Navy program offices and FEL review panel.

A preliminary agenda for the March 10-11 FEL Users meeting and LPC workshop was prepared and is being distributed today to our user email list and posted on the JLab website.

WBS 4 (Injector):

The photocathode gun delivered up to 4 mA of CW beam (350kV, 135pC) for FEL operations during the week. The cathode QE is about 1% after yesterday's FEL lasing success.

Early in the week we did a calibration of the gradients on the quarter cryounit against PARMELA. The code predicts the beam is diverging coming out of the unit for gradients set back in December 1999. However, the actual beam is converging for the present injector configuration.  Assuming an uncertainty of ± 10% in the gradients, PARMELA predicts that the beam  converges if the CAV4 gradient is  10% lower than the Dec. 1999 value.  Later, Steve verified than CAV4 was indeed running at 10% lower gradient than the value read in EPICS. Now the beam behavior (longitudinally and transversely) is closely matched by the code. This is a great step forward in benchmarking our version of PARMELA against the actual injector.

WBS 5 (SRF):

The third cryomodule is complete and will be delivered to the FEL vault on Monday. It will be commissioned alongside the beamline without impacting the 10 kW push in the machine. Once commissioned off line it will be installed at the earliest convenience and should help the drive to 10 kW, if that has not been achieved already.

WBS 6 (RF):

RF - The software for the RF systems was upgraded last night after the machine was shutdown for the installation of the 3rd RF cryomodule.  All working systems were operated into RFON and operated properly.

Installation of 3rd Cryomodule - The back leg of the machine is being opened today to clear a path for the installation of the 3rd cryomodule.  An IOC is being setup so the commissioning of the cryomodule can proceed in parallel with the continued operation of the machine.  The preliminary assembly of the RF waveguides has started, but will not be completed until the cryomodule is located in its interim commissioning position on Monday.  This position is 5' from

its final beam line position.  Download files have been built and are being updated with the external calibration coefficients.

WBS 8 (Instrumentation):

Much effort was put toward preparing systems for the new cryomodule. All parts for the new temperature diode system are on site and being assembled. Cables for the beam line, wave guide, HE vessel and in line diodes were pulled into position as well as cables for the piezo tuners (PZT), and the Higher Order Mode (HOM) loads were installed this morning and are being terminated. DC to DC converters are on order for running the PZTs. A great tutorial on these parts can be found at http://www.piezo.com.  All the RF instrumentation and controls modules were tested and brought on line (thanks to the EES group for their support).  Prepared electronic route for servers from Lab 7 to the various zones. The initial HOM families will be recorded as soon as the module is cold. These will be tracked as the module goes through the commissioning process. 
   Maintenance work has been done on a number of systems and will continue this next week. Last weekend all ~1500 connections to the DC trim magnet power supply patch panels were tightened. There were a number of connections that needed tightening. The main box power supplies will be done on Monday & Tuesday.  Repairs are under way to make the "Coulomb Counters" functional and calibrated. 
   The format was modified for the framegrabbers data files. These provide a way to save all current settings and readbacks for reference at a later date. Worked with Al Grippo, Kevin Beard, and Arne Freyberger to establish a Gaussian fit routine that would require minimal testing for a faster solution. Yesterday an algorithm was tested with successful results. The implementation into the current code is now left with testing to follow. The new machine for the framegrabber was finished this week in preparation for the framegrabber's configuration. Some of the libraries have been installed into the fresh kernel already. 
   The testing of the AMS/Video 32:32 channel crosspoint system (F02553) is complete and finished system will perform well for us. All of the results of the system's throughput performance (bandwidth and phase data) and crosstalk measurements have been archived in the database. This includes the embedded firmware and the list of corrections that need to be made to the prototype. We are now ready to proceed with the production of about 20 boards (F0253) for both the AMS and Video systems upgrades. The documentation for this project is available at:  https://laser.jlab.org/devlore/config.asp?Item=1463 
   Drawings are complete for the front and rear panels for the chassis being used to upgrade to the video system.

WBS 11 (Optics):

This report will be somewhat more lengthy than usual, since we did not file a report last week.  During the week of the 26th, we received new liquid quick disconnects, and verified that they were vacuum tight and compatible.  This will greatly alleviate the torques imposed on our optics from the cooling tubes.  (Kudos to Joe Gubeli and Mike Bevins for thinking of this and expeditiously trying it.)  We also received the stand for the insertable mirror used to route beam into the beam dump in the Optics Control Room.  We also measured the absorbed power due to the THz radiation on our 10.6 um OC mirror.  The agreement with the calorimeter is not great, so more measurements are planned.

Last Saturday we opened the OC vacuum vessel to remove the 90%R (at 6 um) mirror, work on the cooling line arrangement, and determine why we were having problems with the linear actuators.  For the linear actuators, we determined that the problem was due, in part, to an interference with the new cooling ring and the restoring spring that is connected to the LVDT.  We also found some of the spring holders set incorrectly, which adds to the amount of thrust the motors have to provide to move the mirror.  Once these issues were remedied the linear actuators worked fine.  The mirror removal was done to try cryocooling it this week.  However, we found that it was loose in its holder, so we have spent this week recoating the edge and remounting it.  It will be ready for testing by Monday, then installed in the optical cavity for higher power lasing.  We also determined that its reflectivity curve was unchanged since it was new, so that was not the reason it wasn't lasing.  Tuesday, we aligned the 98%R (at 6 um) mirror, and were gratified to see that our work had resulted in a considerably improved mirror figure.  We then lased on that mirror late in the day.  By Thursday evening (when we had a Brewster window crack) we had produced 1.51 kW of output.  While operating at this power we recorded the outlet water temperature and flow rates and were able to make some determinations of the losses in both optics.  The loss in the outcoupler, assuming a 5 W/mA contribution from the THz, is

~ 550 ppm.  Somewhat surprisingly, the HR had a higher loss of 3700 ppm.  Even then, we still had half our control margin for the deformable mirror.  Thus, we feel we have a good shot at high power once we reinstall the 90%R mirror.  This is planned for early next week.  We also made excellent progress in constructing a vacuum calorimeter that will be placed in the vault next to the diagnostics hutch.  We'll use it to determine the loss in optics we're receiving from vendors, and from our IR Demo mirror set.  We also reinitiated efforts to look at other bonding techniques for our mirrors.

We held planning meetings to scope the level of effort to implement the recommendations of the Semi-Annual Review Committee.  Specifically, meetings were held on the cryocooled ZnSe effort and the scraper output.  A supporting document on the cost and delivery of optics to support various wiggler options was prepared.  This week the FEL equipment and staff endured a lot of microphonics and noise as the penetrations for the UVFEL optical transport were drilled in User Lab 1 (which serves as the local office for half the Optics Group - sorry guys, it's almost over!).  Another important milestone needed to get the UVFEL brought to fruition.

Terahertz Project:

The bids for the mirrors were received.  In view of the large minor axes of the ellipses, a free-form machine is necessary for manufacture.  Only 2 vendors could supply this, and only one vendor could meet our delivery specifications.  The first mirror is scheduled for delivery on March 20. 

The final order was placed for the mirrors.  Vacuum components are starting to arrive and are being consolidated in Lab. 4.