# MEMORANDUM

To:		Distribution

From:		F. Dylla

Subject:	FEL Upgrade Project Weekly Brief  - June 13-17, 2005

Date:		June 17, 2005

Highlights:

We devoted most of this week to tweaking the machine for high power operation at

2.8 microns.  By week’s end we had optimized the lasing efficiency at low currents:

lasing efficiencies exceeded 1.5 kW/mA  at low pek charge (70 pC) and we obtained 2.5 kW  cw at 2.5  mA  and 3.5 kW at 5 mA.   Metrics of excellent lasing were observed: long detuning curves (12 microns); short laser turn-on times  (8 microseconds).  We obtained a good view of where we are and we can come back to optimizing this configuration after our planned two week shut down from June 20-July 5th.

During this shut-down we will be installing the high power 1.6 micron optics, installing and replacing various beam viewer diagnostics and vacuum isolation valves, and laying out a good portion of the UV FEL chicane.

Also this week, we completed the laser certification of the experimental hutch in the THz User Lab (3a) and obtained THz spectra with this new configuration.

The first of the 750 MHz cavities for the 100 mA injector which were manufactured by AES, completed processing and was tested at 2K.  The measured gradients (28 MV/m) exceeded specifications by a factor of two. This result equals the best of the performance we saw with the entire production run for 804 MHz cavities for the SNS project.  These results are timely for today’s ONR/JTO project meeting held in Arlington for the AES/JLab 100 mA DC injector project.

Tonight before we shut-down for the maintenance/installation period we will be qualifying the transmission of high power 2.8 micron light into User Lab 1 and delivering some test beam to the NASA/CWM nanotube experiment.

Management:

The project monthly financial reports for May were completed and forwarded to the Program Offices.  Detailed planning of the staffing and budget allocations for FY2006 continued.

On Wednesday, the detailed plan for next week’s shut down was reviewed. 

This week had a number of interesting and productive meetings with visitors.

On Monday and Tuesday of this week we hosted Peter Weightman, one of the 3 principal investigators on the 4GLS (FEL) project in the UK, and 3 colleagues. Their purpose of visit was to investigate opportunities for scientific collaboration in the THz regime.  They have received new funding to establish a THz beamline and facility and they presented a preliminary design.  However, they expect several programmatic interruptions due to scheduling of the first UK FEL (the ERLP), and the upgraded machine. They would like to bridge those periods by collaboration with JLab.

On Wednesday we hosted Prof. Dave Reitze from the Dept. of Physics at the University of Florida. Dave is PI with the LIGO Project who is involved with the high power optics and lasers which are the critical components for these two large gravity wave observatories.  We have been keeping track of LIGO’s progress in the development of low loss, high power optics at 1 micron.  Their current state-of-the-art is optics that have absorption losses less than 1 ppm and scattering losses less than 20ppm.

On Thursday, we hosted about a dozen scientists and collaborators from the US Army Night Vision Lab (Electronic Services Division).  At their request we discussed the capabilities of the FEL and THz laboratories and the visitors toured the facilities.  This meeting will help them to formulate and co-ordinate future projects utilizing our unique light sources. The Army NVL funds our THz program in collaboration with AES, Inc.

Also on Thursday we hosted a visit by Prof. Avi Gover, from the Physical Electronics Dept. at Tel-Aviv University, who also directs Tel-Aviv’s FEL Facility. Prof. Gover gave a seminar on recent results and studies from his group.

We were informed that the FEL project proposed to the NSF from Florida State received very favorable reviews in the first go-around.  The NSF asked for clarification of some of the points raised by some of the referees, and this week George Neil and Gwyn Williams participated in helping to address these.  FSU has asked JLab to be a partner in this proposal.

Operations:

It was an interesting week.  Operations concentrated on getting a good high power setup at

2.8 microns, 88 MeV.  We struggled this week to get both good matches to the wiggler and good matches from the wiggler to the dump.  We succeeded, perhaps a bit too well, in the first and only partially succeeded in the second. 
   The first part of the week was spent in careful setup of the  accelerator.  Betatron matches were found, the dispersion was suppressed, and the longitudinal match was optimized.  The laser  lased reasonably well with this injected beam. 
   We had not operated at high current yet with the latest injector setup with stronger solenoid focusing operated 5 degrees off crest in cryounit cavity 4.  When we did, we found that the pressure rose at the exit of the cryounit proportional to the current.  Carlos Hernandez-Garcia looked at this in PARMELA and found that, indeed, the beam was too large coming off of the unit.  At this low an energy you need more focusing at the exit of the unit.  The best way to get  this in our injector is to use the first cryounit cavity (cavity 4)  focusing.  This does reduce losses to reasonable levels but increases the bunch length going into the linac.  This increases the  energy spread more than we would like.  We were able to get a good match with this beam to the wiggler in both the transverse and  longitudinal phase spaces.  The result was very strong lasing.  The laser turned on to half power at the peak of the gain curve in about  8 microseconds.  This is much faster than we have seen to date and implies a gain of about 80%/pass.  The cavity length detuning curve is also very long, also implying high gain.  The exhaust energy spread, unfortunately was huge.  It is at least 12% full width.  This is sufficiently large that energy recovery is nearly impossible.   Half of Thursday was spent in a futile attempt to reduce losses  sufficiently that we could operate at the peak of the detuning  curve.  On Friday we tried to improve matters by reducing the charge.  The interesting thing is that we were able to achieve 1.5 kW/mA efficiency with only 70 pC of charge.  This corresponds to an efficiency of 1.7%.  Since not all of the charge is participating, this implies that the part that is lasing has about 2% efficiency.   One expects an energy spread from 12% to 16% for this efficiency.   When an efficiency of about 1.5% is exceeded the beam blows up near the dump in the last cryomodule and trips off one of the last cavities.  This is due to uncorrected chromatic aberrations for such a large energy spread.   With a little tuning we were able to get 2.5 kW with 2.5 mA of beam and 3.5 kW with 5 mA of beam at low charge.   More careful optimization will be needed to do much better than this. 
All this argues for several things: 
1  A cryomodule with a 3 inch aperture in the third position would be nice (of course this is guaranteed in lower frequency structures.)

2. Larger apertures in the injector would also be nice.

3. A higher injection energy is required for this design.  We did look briefly at operation with

8 MV from the injector instead of the 7 MV we have now.  We found that we should be able to run there without too much field emission induced vacuum rise.  We can look at this when we come back up. 
   Other activities were conducted in parallel with the power push.  An  effort was launched to align the streak camera to the light emitted  from the first arc.  Some more work is needed on this.  User lab 3  was certified for THz operation.  We can now do THz spectra and make  modifications to the spectrometer as we wish.

We exhumed an old design model used back in '01 to lay out the 4F and 5F energy recovery matching telescopes. This model fits beam envelopes from wiggler to linac across the arc, through a range of momenta. Though only crudely approximate to the installed machine (it lacks, for example, the THz management chicane), it gave hints as to some of the issues we are facing. First, the level of chromatic aberration in the initial optics solution was pretty certainly seriously bad. Secondly, the spots (with bright core and halo) we saw at 4F09, 4F12, 5F05, and so on can be understood as showing the on-momentum betatron core and the chromatically mismatched tails. Third, the wretched transmission down the linac on Thursday was due in part to vast input momentum spread from the insanely highly performing gonzo-journalistic FEL talking to a pair of telescopes that had/have not been properly chromatically balanced. Fourth, the improvements seen on Friday using a meatballed match correspond to a better balance of the telescopes. Finally, and most tantalizingly given the impending shutdown - there appears to be room for additional improvement.

WBS 4 (Injector):

The DC photocathode gun is operational delivering beam for FEL ops with a 43% request on the drive laser polarizer. A QE measurement was performed this week by measuring the drive laser power needed to run 1 mA cw. The QE was about 4%. 
   The aluminum cylinder for the NEG coating solenoid was submitted to the machine shop for fabrication. The associated solenoid coil assembly drawing was completed. Polishing on the segmented support tube continued. More equipment from the FEL lab was moved to the ARC building lab. 
   The Si02N coated electrode was conditioned further and showed stable operation at 25 MV/m, higher gradients generated higher emission at acceptable levels. 
   Work continued this week on benchmarking the 100 mA injector model with AES results using  PARMELA LANL version.

WBS 5 (SRF):

This week we made a small but significant step by getting first hardware in the ampere-class cryomodule project. It may not look like much but from little acorns mighty oak trees grow, or so we hope. These are parts for a copper model intended to validate the calculations of the waveguide HOM dampers. We are doing this at 1500 MHz so we can test them on a CEBAF single-cell cavity and get results quickly. The waveguide dampers are key to obtaining BBU thresholds in the Ampere range while keeping the cryomodule as short as possible. We are also making steady progress on the 750 MHz cavity concept, tuner concept, cryomodule layout and other ancillaries. The next step after the copper model test is to make a niobium prototype.

WBS 8 (Instrumentation):

The Terahertz lab was also completed (HONEST TO GOD!) this week. After doing a little early morning searching in the lab and racks, the camera for inside the hutch is in the system. There is also another camera outside the lab that will be hooked up by the end of the day. The hutch/room were then certified and commissioned to run terahertz beam into the lab. The software for the power meter on the FTIR was also installed on the new(er) dell inside the lab. After giving us some fuss, the power program was cleanly installed on the PC and seems to be working well.
    Support for the video system update began this week with the running of sixteen RG-59 cables from the video switcher rack to the control room. The added features will be 32 instead of 16 outputs and up to 256 inputs. We are ready! All aspects of the new video system have been fully tested and are ready to be installed at the start of the maintenance down. We have 6 chassis on the bench which have been tested and are ready to be installed. All of the CAN bus wiring has been tested and each chassis has had its unique CAN system ID set. We have also relabeled every cable in the system to include a unique database ID. In the process, this act has functionally documented our existing channel set-up and will allow us to quickly find "the other end of that cable" and the full history of what it has been used for and what it has been connected to as we proceed from here with our video system configuration.
    A final configuration and design for the 6kW polarity reversing switch has been completed.  EECAD is currently working on the schematic for the system PCB.  A prototype enclosure has been modified to incorporate all the necessary connections and include the switching electronics.  The power, control, and status will all come from the same location so an external power source will not be needed.  A channel flow chart is being completed to provide a graphical explanation of how the system will function.  The backplane design has been completed for the beamviewer crate upgrade.  The firmware programming first draft has been completed for the beamviewer control card, this will provide the local bus control for each card in the crate.  The design for the card is still in progress with several details still being worked out.
    Activities are also coming to a close in Lab 4. The security camera is correctly in the system along with the area monitor inside the clean room in the lab (after some serious snooping to find the power cables for them). The two outside monitors are also in place and just awaiting channels in the new video upgrade that will take place. Two of the LPSS boxes were hung this week in lab 3A and lab 4 after the shipment of unistrut and brackets arrived for them. The enclosure for the LPSS Master that was sent to the machine shop also arrived and is ready to be assembled. 
    Two of the revision C Cryo-Diode thermometer dedicated F0273 boards came in assembled and were tested successfully. One of the two new boards was installed in the already installed CTD chassis to fill it out (16 total boards, 64 channels). 
    Extensive testing of the circuit boards for the GC power supply has taken place. The boards were tested with the magnet setup both individually and in pairs to determine the best combinations available.  Screen shots of possible oscillations were taken, and from these the source of the oscillations should be found. 
    Another beam viewer chassis has been repaired and checked out, and is ready to support the UV upgrade.  The beam viewer test boxes have also been repaired and are ready for use in aligning the new viewers. Testing of the 4 Channel Solenoid Fail-safe Interlock PCB and box has been completed.  The system is ready for installation during next week's scheduled shutdown.  A 4 Channel Limit Switch Simulator box was also designed, built, and tested in order to make testing of the system less complicated.  A debugging procedure is being prepared for the hardware. On a side note, to prepare for the upcoming shut down, we will give a thorough check of all the completed CCD beam viewer cameras to make 100% sure the ones that are tagged do and will work. 
    Moving in the new trailers continues; another shelf has been installed over there, and all of the power supplies and transformers have been relocated there.  Additional shelves have been assembled and stored in the tent in continual preparation for the trailer move in. 
   The needed correction to the lab safety application was made to allow for the final certification of lab 3 (THz lab). The IOC reboot to install the lab safety correction caused the polarizer to become lost, creating a problem with the machine's charge per bunch. The polarizer position was corrected with the help of Shukui Zhang. 
    We sent Richard Walker the miniphase procedure to help formalize it. Working with David Douglas to record data on the ITV0F04 viewer to form a reference point for after it is changed out during the upcoming down. The automated miniphase (with the help of Matthew Stokes) is coming along. The schematics for the required EPICS records have been drawn up but not installed. Created an EPICS record that maintains a rolling average of a given sample size (this will be used by the CEBAF software staff in the future). We are working with Steve Benson for the design of the goggle safety system. This will be an addition to the Laser Personnel Safety System. Received HID Corp RF tag options for the goggles. Looking into an Omron solution that will better mesh with our LPSS' PLC configuration.
     Renascence cryomodule assembly continues; Terminated Piezo, Motor Drives, Limit Switch wiring through the tuner ports on top of module. We will install connectors once they are leak checked.  Next week we will be ready to install the 8 top hats, then we can terminate the Vessel diodes, heaters, probes.
       Coldfire 5282 mirocontroller was set as Autoboot, so that the bootloader would pass control to the RTEMS OS image after 10 seconds. The RTEMS will download database and records from an NFS server and initialize the IOC. Database, device supports, and driver for beamviewer control are under developing. Awaiting the beamviewer control card for testing. 
     We are building up another "mirror raster control board" for the "photocathode response mapping system" has been a major activity of the week. We now have three of these boards built-up (including the one installed) and we have documented the three jumper-wire modifications required for the board to work. At the beginning of the down, when they map the cathode, we plan to swap out the existing board and do an in-system test of our new spares.

WBS 9 (Beam Transport):

Sextupole (SF)

•	Work continues at New England Techni-Coil. Cores were assembled and marked. Secondary taps to individual pieces is started. Pinning of cores starts late next week. First coil potting had the inner mandrel stick to the coil. They are working the issue. A vendor for ceramic tubes that act to isolate the LCW circuit from the coil was found by Jlab’s Larry Munk and Tim Siggins.

Replacement Chicane Dipoles (GW)

•	Work continues at New England Techni-Coil. The third core gets Purcell gap glue-down starting Monday. The seventh coil (of 8) is almost wound. Potting of the third coil starts next week. All core machining is complete.

UV Line

• 	Set ability of the QX Quadrupoles of the UV Line was the object of our investigations using the hall probe in one magnet this week. Magnet Test found the Bang-Bang Hysteresis and Setting Protocol was just as good as the Slow Ramp and has a good likelihood of repeating in the FEL enclosure. Epics and trim card control of the field during tests can now be done. Set ability of the actual field gradient integral is still compromised by drift in the rotating coil stand at the level of ±2.5 parts in 1000.  Ken Bagget is running tests using different data gathering to attempt to reduce this band of uncertainty. In the mean time, they believe they should be able to compare the strength of all quadrupoles to the “Standard” Quad that was measured a SLAC using a bucked measurement using the existing, long 2 inch probe.

• 	Girder assembly drawings and drawings for piece parts for Region 6Fsub1 were completed and are now in checking. Regional Assembly of the 8Fsub4 (Last) is in preparation

• 	Larry Munk is working at getting the existing inventory of surplus girders converted to the required girders. We have seven that we can start assembly with on July 6.

UV Wiggler Progress

•	No progress on the new Wiggler vacuum chamber.

IR PM Wiggler Vacuum Chamber Progress

•	The design of the wiggler chamber started.

WBS 10 (Wiggler):

STI Optronics has received the support structure for the IR wiggler.   They expect the guide rails next Tuesday.  They tested the limit switch repeatability and found that it far exceeded our  specification.  They also looked at noise in the Heidenhain gauges and found none even when all four motors were being driven. 
   We had the kick-off review for the UV wiggler.  Since STI Optronics is using most of the design from the IR wiggler for the UV wiggler, the review was pretty simple.  A few minor action items were found.   The design overall is very nice and should provide excellent service  for the UV device.

WBS 11 (Optics):

FEL mirrors:

This week we completed the debonding of the 1.6 micron OC and rechecked it's figure.  We now have a good optic.  The 1.6 micron HRs arrived, and have a longer ROC than we'd like.  We believe we found the problem with the astigmatism in the 1 micron mirror; it is due to the silicon dioxide coating on the back face of the optic.  While it is a rough surface, we don't have good adherence.  We've spoken to the coating and glue vendors, and are regluing the optic after trying the quickest of the recommendations, to soak the coated surface in acetone for 1 hour.  The

1.06 micron, 96% R OCs will ship next week.

Other activities:
   We continue to juggle a number of activities this week.  We have been moving all hardware not required for the next 6 months buildup activities.  (A thanks to the gun group for ratcheting up their moving activities to make more space for us.)  A drawing for a sizeable hutch for the Dahlgren group was produced and sent to them for comment.  We worked late last Friday to obtain more beam profile data in their hutch.  (Thanks to Chris Behre for working late with them.)   We have continued work on refurbishing mirror cassettes for Labs 2, 4 & 6.  We ran high average power into User Lab 1 to look at the transmission.  We have reasonable transmission into the OCR, but low transmission into the lab.  We answered questions from staff from Daresbury about our procedures. We also met with Prof. David Reitze, Univ. of Florida.  He is associated with the LIGO optics core group measurements, and was very helpful in providing corroborative information that verifies that the loss of their optics (at 1064 nm) is about 1 ppm.  We welcome Jason Taylor, a student from Old Dominion University, to the team for the summer.

UV FEL
   We are working through the list of comments received from the optical cavity review.

Terahertz Project:

The hutch in lab 3 was certified.  Visible synchrotron radiation was observed on both diamond windows.  The TK Associates THz absolute power meter interface card was installed and the meter was set up in lab 3. Beam was restored to the Michelson and spectra were taken.