Tune Mode Generator Setup:

This note outlines a procedure for installing and aligning elements for producing a waveform (on/off) of optical pulses, our so-called Tune Mode Generator, from a CW laser.

Equipment needed:

- Laser (1)

- Glan-Thompson Prism Polarizer (2)

- Halfwave plate nominally close to the laser wavelength (1)

- Pockels cell (1)

- Spinning analyzer (1)

- Power meter (1)

- High voltage Pockels cell driver capable of  (QWV (quarterwave voltage) at 50% duty factor and either polarity HWV (halfwave voltage) at desired waveform. (1)

- 2-channel oscilloscope (1)

Procedure:

- About 12”-18” are required for the optics and alignment diagnostics.

- Layout the elements in the following pattern:

L1 : P1 : HWP : PC : SA : P2 : PM

L1 = laser

P1 = cleanup polarizer

HWP = halfwaveplate

PC = Pockels cell

SA = spinning analyzer

P2 = pulsed mode polarizer

PM = power meter

- Insert L1 and PM.  Measure the laser power _________ at PM.

- Insert P1 and align P1 _________ for max power _________ at PM.

- Insert P2 and align P2 _________ for max power _________ at PM.

- Insert HWP and align HWP _______________ for min power _________ at PM.

- Retract HWP, providing it can be equally re-inserted later (e.g., via controller).

- Insert PC, SA, and align using standard Pockels cell alignment procedure using (QWV at 50% duty factor.

- Remove SA and re-insert HWP.

- Suitably connect PM output to oscilloscope and apply HV waveform to PC.

- Adjust HV driver HWV (max. signal) and Null (min. signal) to achieve desired waveform on oscilloscope.  It is helpful to trigger the scope on the waveform.

- To help achieveing a good min. signal it may be necessary to adjust be a couple degrees the angular orientation of the HWP to “bias” the elliptical polarizartion slightly in the opposite direction on which the HV driver Null setpoint acts.

- Remove PM.