## Slide 1: Soft X-ray Self-Seeding
Philip Heimann (SLAC)
Daniele Cocco, Juhao Wu, Jim Welch, Yiping Feng, John Amann, Zhirong Huang, Jerry Hastings (SLAC)
Paul Emma (LBNL)
SLAC/LBNL R&D project in Soft X-ray Self Seeding

## Slide 2: Soft X-ray Self-Seeding Concept
SASE FEL x-rays are generated in a 1st undulator section.
A grating monochromator selects a narrow x-ray bandwidth.
The electron beam passes to the side in a chicane.
The x-rays from the monochromator seed the FEL x-ray generation in a 2nd undulator section.
Proposed by J. Feldhaus, E.L. Saldin, J.R. Schneider, E.A. Schneidmiller, M.V. Yurkov, Opt. Comm., V.140, p.341 (1997)
  Not implemented at FLASH
2nd undulator
M1
M3
G
h/2
g’/2
M2
e-beam
Source
plane
Re-entrant
plane
1st undulator
S
x/2

## Slide 3: Motivation
SASE FEL pulse is longitudinally incoherent
Soft x-ray self-seeding
  Reduce spectral bandwidth
  Remove spectral jitter
  Make a near-Gaussian pulse in time
SASE FEL longitudinal profile at 26 m
SASE FEL temporal profile

## Slide 4: Symmetric Design (Toroidal grating)
1.2 m
0,0
60 max,3.85
1350,3.85
1535,3.85
1663, 0
~1290 mm
Fit within the length of one undulator module, 4.5 m.
Photon energy range 400 - 1000 eV.
X-ray and electron delay varies from 660 - 850 fs.

## Slide 5: Beam Transverse position @ midpoint of chicane
8 mm
16.1 - 19.9
3.85
0
x-ray
electron
X-ray and electron deflections are in the horizontal plane.

## Slide 6: Symmetric Design (Toroidal grating)
Central groove density (l/mm) | 1123
D1 (l/mm2) | 1.60000
Radius of curvature (m) | 195
Diffraction order | 1
Fixed incidence angle (deg) | 89.0000
Sag Radius of curvature | 18 cm
Resolving power from 7800 (400 eV) to 4800 (1000 eV).
D. Cocco

## Slide 7: Pulse stretching vs resolving power
Grating x-ray pulse stretching Dt =N m λ / c.
The grating x-ray pulse stretching 1.7 times transform limit.
X-ray pulse will be longer than electron bunch.

## Slide 8: Beam steering
slit
spherical
-0.5 mm
+0.5 mm
plane
M2 15 mrad
Incidence
Overlap of x-ray and electron beams made by translation or rotation of M2 and M3 mirror.
M3 15 mrad
Incidence

## Slide 9: Overlap scheme
U9
U10
U11
YAG
YAG
Use x-ray steering (x, x’, y, y’) to move x-ray spots on top of electron spots on both Ce-YAGscreens.
SXRSS
12 m
σ≈35μm
σ≈35μm
U8

## Slide 10: Transmission
Chart: 
| 400 | 450 | 500.00099999999969 | 550 | 600.00099999999998 | 650 | 700.00099999999998 | 750.00099999999998 | 800.00199999999938 | 849.99900000000002 | 900.00199999999938 | 949.99800000000005 | 1000
Monochromatic Transmission | 4.6669 | 5.0031 | 5.176 | 5.1152 | 5.2477 | 5.3282 | 5.3447 | 5.2919 | 5.3433 | 5.4054 | 5.4108 | 5.3359 | 5.1535
Including resolution and with 0.3% SASE bandwidth.
Laminar profile
w
h
Pt optical coatings

## Slide 11: Spot expected in the following undulators
Distance from M3 | Horizontal spot size (mm) at 400/1000 eV | Vertical spot size (mm) at 400/1000 eV
2 m | 67/66 | 40/32
Based on geometric ray tracing.
Future work coherent beam propagation.

## Slide 12
Parameters and longitudinal phase space area after Gaussian fit to both temporal and spectrum distribution are summarized as follows (defined as stsw)
Seems to be 2 ~ 3 times of transform limited
Cases studied and results
undulator | 1.2 nm (1 keV) |  | 2.5 nm (500 eV) | 
 | 20 pC | 100 pC | 20 pC | 100 pC
LCLS | 0.77 | 1.18 |  | 
LCLS-II | 0.94 | 0.98 | 1.69 | 1.48

## Slide 13: High peak power
@ 60 m
1 keV Soft X-ray Self-seeding (10 kW after mono)+ Taper  350 GW
~ 100 pC Gaussian temporal dist.
Longitudinal phase space: ~ 2 times of transform limited
@ 60 m
U33
2.2 x 10-4
fwhm
Grating monochromator

## Slide 14
Summary
At the LCLS soft x-ray self-seeding is possible in the length of one undulator module.
The optical-electron design is nearly complete.
This project is a collaboration between SLAC and LBNL.