## Slide 1: Practical Aspects of SRF Cavity Testing and Operations
Tom Powers
SRF Workshop 2011
Tutorial Session

## Slide 2: PRACTICAL ASPECTS OF OPERATION
Interlocks
Optimizing Loaded-Q
Pulsed response as a function of loaded-Q
Multiple cavities driven by a single source

## Slide 3: CRYOMODULE INTERLOCKS
NEVER OPERATE A CRYOMODULE WITH THE COUPLER INTERLOCKS BYPASSED
Coupler Interlocks
  Arc detector(s)
  Coupler vacuum
  Window temperature
  Water flow (If water cooled)
  Electron probe (Useful but not required)
  Water temperature (Useful but not required)
Cryomodule
  Cavity vacuum
  Helium level
  Helium pressure (Useful but not required depending on cryo plant)
  Insulating vacuum (Useful but not required)
RF Driven Interlocks
  Quench detection
  E2/PFWD ratio
  Gradient Present with RF off

## Slide 4: OPTIMIZING LOADED-Q
Where
  δf is the frequency shift of the cavity from the generator frequency
  ψB is the phase of the resultant* beam relative to cavity gradient and
  β = (Q0 – QL) / QL
You need to take the derivative of this equation with respect to QL in order to calculate the minimum klystron power necessary.
* Note: For an multiple beams at once, i.e. an energy recovered linac (ERL), the resultant beam current is the vector sum of the beams

## Slide 5: MODERATE AMOUNT OF MATH

## Slide 6: REDUCED SOLUTION FOR SRF CAVITIES OPERATED ON CREST
But life is never perfect

## Slide 7: THE EFFECTS OF TUNING ON OFF CREST CW BEAM LOADING
On beam turn on the forward power increases the phase shifts and microphonics effects are multiplied
The tuner operates with a goal of making ψKly equal to zero by shifting the frequency by δfS which compensates for the I0RCsinψB term.
Thus ψKly → 0 and PKly is minimized to:
Where δfM is the frequency shifts due to microphonics
Thus in this case:
0

## Slide 8: THEORETICAL EXAMPLE OF TUNERS COMPENSATING FOR OFF CREST BEAM LOADING
Chart: Time (Sec)
CEBAF 7-Cell Cavity
  L = 0.7 m
  (r/Q) = 960
  E = 8 MV/m
  QL = 2 x 107
  δf = 10 Hz
  I0 = 10 mA
  Pass one ψB = -10º
  Pass two ψB = 166º
  Resultant Beam
  0.7 mA at 78°

## Slide 9: THEORETICAL EXAMPLE OF TUNERS COMPENSATING FOR OFF CREST BEAM LOADING
Chart: Time (Sec)

## Slide 10: PREDICTED AND MEASURED FORWARD POWER IN AN ERL
Chart: Beam Current (mA)
The solid lines indicate the predicted values based on:
  QL = 2 x 107
  E = 5.6 MV/m.
  Δf = 10 Hz
Test Process:
  Tune the cavity with no current.
  Disable the mechanical tuners.
  Ramp the current up and record the forward power and phase.
  Repeat with Tuners enabled.

## Slide 11: Predicted and Measured RF Drive Phase In an ERL
Chart: Beam Current (mA)

## Slide 12: EFFECTS OF MICROPHONICS AND IMPERFECT ENERGY RECOVERY IN AN ERL CAVITY E = 20 MV/m, I0 = 100 mA

## Slide 13: SELECTING LOADED-Q FOR OFF CREST BEAM
Selection of loaded-Q has implications on RF power requirements.
When the beam is operated on crest the process is straight forward and margins only have to be added for
  Microphonics,
  Uncertainties in cavity parameters such as QL and operating gradient.
  Overall Margin
  Detuning effects.
When the beam is not operated on crest operational modes must be considered. Often this can substantially reduce the RF power requirements.
  Ramping current simultaneous with operating tuners.
  Allowed levels of pulsed operation
  Uncertainty of the relative beam phases in an ERL

## Slide 14: MULTIPLE CAVITIES ON SINGLE SOURCE
It can be a desirable to use a single source to drive multiple cavities.
  Reduced cost per Watt at higher RF –Power levels.
  Availability of klystrons or IOTs at desired levels for multiple cavities.
  Unavailability of klystrons, IOTs at desired power levels for single cavities.
  Reduced number of LLRF systems to drive cavities.

## Slide 15: MULTIPLE CAVITIES ON SINGLE SOURCE
It can work fine when:
  The cavities are operated near crest.
  The beam is not sensitive to minor variations in gradient and phase.
  Loaded-Qs are well matched.
  Gradients are close to the same for all cavities.
  The loaded-Qs are relatively low as compared to pressure sensitivity and microphonics.
  You have the advantage of a large number of cavities and individual errors are corrected by statistics.
It can present problems when:
  The beam is sensitive to errors in gradients or phase.
  Detuning becomes significant as compared to the FPC bandwidths.
  Cavities are operated at different gradients.
  Cavities have different loaded-Qs
  Cavities are operated at different beam phases with respect to crest.
  While linacs are an area where this concept can be very practical, injectors are an area where the problems become important especially when space charge and cavity induced beam focusing are important.

## Slide 16: SIMULATION METHOD
Use the basic complex RF voltage to complex gradient equation to calculate the field in each cavity, including beam phase and cavity detune angle.
Sum the real and imaginary parts of the electric field.
Compare the vector sum to the desired vector sum and calculate the error in the vector sum.
Add, with gain, the complex error to the complex RF voltage from the current pass.
Use this sum to calculate gradient in each cavity.
Repeat until the real and imaginary parts of the vector sum error are below a threshold.
Where ψ is the cavity detune angle.

## Slide 17: ERROR IN GRADIENT AND PHASE WHEN 1 OF 3 CAVITIES IS DETUNED
QL = 2e6
F0 = 1497 MHz
(r/Q) = 960 Ω/m
L = 0.5 m
β = 3000

## Slide 18: PHASE ERROR IN WHEN THE PHASE IN 1 OF 3 CAVITIES IS DIFFERENT

## Slide 19: GRADIENT ERROR IN WHEN THE PHASE IN 1 OF 3 CAVITIES IS DIFFERENT

## Slide 20: Gradient Error in When the Loaded-Q of 1 of 3 Cavities is Higher Than the Others

## Slide 21: CONCLUSIONS
Thank you for your attention. I hope what I have presented will be useful.
Selected References.
Padamsee, Knobloch, and Hays, RF Superconductivity for Accelerators, John Wiley & Sons 1998.
T. Powers, “Theory and Practice of Cavity Test Systems” 2005 SRF Workshop. (Note: Includes all of the math for testing cavities)
T. Powers, “RF Controls Experience with the JLAB IR Upgrade FEL”, 2009 ERL Workshop.
J. Delayen, et. al., “Development of a Digital Self Excited Loop for Field Control in High-Q Superconducting Cavities, SRF 2007.
L. Merminga, “RF Cavity Equations Steady State”, JLAB Technical note TN-95-019.
K. Davis “Microphonics Testing of the CEBAF Upgrade 7-Cell Cavity”. PAC 2001.