## Slide 1: Hall A Tritium Target Status
Dave Meekins
18 January 2017

## Slide 2: Overview
Intro
Subsystem Status
  Cells
  Shipping
  Exhaust
  Beamline
  Cryo-System
  Controls/Safety
  Administration
Summary

## Slide 3: Information and Documentation
New wiki for the tritium target:
  https://wiki.jlab.org/jlab_tritium_target_wiki/index.php/Main_Page
  Reports: design and technical
  Drawings
  Procedures
  Calculations
  Presentations
  Review
    Presentations, Responses, Committee reports

## Slide 4: Target Ladder
Heat sink cooled by ESR to 40K
Stabilized by heater
5 Cells:
  T2,H2,D2,He3,MT
Optics
Dummy/alignment
Solid
  BeO, C-hole, C,

## Slide 5: Tritium Target Final Ladder
Target | Total Solid Thickness | Fluid Thickness
Tritium Cell | 0.16 g/cm2 | 0.075 g/cm2
Helium 3 Cell | 0.16 g/cm2 | 0.075 g/cm2
Deuterium Cell | 0.16 g/cm2 | 0.12 g/cm2
Empty Cell | 0.16 g/cm2 | N/A
Hydrogen Cell | 0.16 g/cm2 | 0.05 g/cm2
Aluminum Dummy | 0.3 g/cm2 | N/A
Carbon Optics | 0.2 g/cm2 | N/A
Carbon Foil | 0.167 g/cm2 | N/A
Carbon Hole | 0.167 g/cm2 | N/A
Thick Aluminum | 1.37g/cm2 | N/A
BeO | 0.16 g/cm2 | N/A
Be Isolation window | 0.04 g/cm2 | N/A

## Slide 6: Target Cell
Main Body and Entrance Window
ASTM B209 AL 7075-T651
Valve assy:
SST 316 and 304
1090 Ci of T2 (0.1 g)
~200 psi at 295K
25 cm long
ID of 12.7mm
Volume = 34 cc
Aluminum CF seals

## Slide 7: Cell Cross Section
0.018” wall

## Slide 8: Operational Limits
Beam Current
  25 µA on all others
Raster
  2x2 mm minimum
Beam Trip limit (18000 trips)
Temperature
  Heat sink must be ~40K
Target Operator 24/7 when cold
Tritium effects access procedures and training requirements

## Slide 9: Tritium Compatibility Study
Tritium and Hydrogen Compatibility
  Extensive experience with H2
  Beam induced corrosion not expected below 180K (Flower et. al.)
  Beam assisted embrittlement
    Increases fugacity
      Atomic tritium recombines rapidly
    Modeled room fugacity ~3100 psi
    Below threshold for H2 embrittlement
    Many orders of magnitude below He-3 swelling threshold (Louthan)
Test SRNL/SRTE using precracked coupons exposed to 2500 psi T2 complete
  Test follows ASTM 1820 G168 (Precracked Stress Corrosion Testing)
  Samples exposed for 4, 8, 12 months
  M. Morgan, A. Duncan (SRNL)
  No measurable effects observed

## Slide 10: Current Status
Six acceptable cells have been fabricated
One complete assembly (including shipping covers etc.) is at SRS
  Used for training and mockup
One cell to be used for Argon experiment in Spring
Fabricating 2 more
Cell
  Accepted by JLAB review
    ERR Sept 2015
  Accepted by SRS review
    Safety Basis October 2016
Status (100% Complete)
  We have enough to run with no spares

## Slide 11: Exhaust System
36” OD 20m tall Stack
12000 cfm blower multispeed
2” pump exhaust
  Run parallel to stack
Stack must also serve as smoke removal
Provides controlled release of secondary and tertiary containment
Pump exhaust is continuous
Blower activated:
  Manual
  Interlocks

## Slide 12: Exhaust Routing

## Slide 13: Stack Base and Blower

## Slide 14: Stack lift/assembly

## Slide 15: Internal Exhaust

## Slide 16: External Components
Stack (Complete)
  50 ft above Hall A Dome access (20 m above grade at site boundary)
  Applicable Codes and Standards
    ASME STS-1 (2011)
    ASCE-7 (2010)
    UL-96
    JLAB SE approval
Blower (90% Complete)
  12000 SCFM blower variable speed (VFD)
  Applicable Standards
    JLAB FPE Approval
  Not connected to Hall port at this time
Pump Exhaust (80% Complete)
  Vacuum pump exhaust parallel to stack
  Not connected to Hall port at this time

## Slide 17: Internal Components
Handling Hut (100% Complete)
Platform (90% complete)
  Needs fitting
  Installed only for Cell installation/removal
Ductwork/Damper (99% Complete)
Truck Ramp Door Control (99% Complete)
  Limited to very small number of individuals
Cans Access (100% Complete)
Fan Controls (75% Complete)

## Slide 18: Beamline Alterations
No plans to substantially alter beamline
Upstream beamline shall be isolated by a Be window
  0.008” thick 1” ID.
  Water cooled (3W beam power 25 µA)
  Reentrant (Resides in chamber)
Downstream beamline to be isolated if effects to the physics are not unacceptable
Window is 15 cm from entrance to cell
Densimet collimator 10 cm long installed in tube upstream of window. (W 90% , Cu 8%, Ni 2%)
Maintenance is possible if required.
12 mm thick collimator attached to cells
Collimators should prevent steering error from affecting cell
  Last steering element is 8 m upstream and 2” radius beam pipe.
Status (30% Complete)
  Collimators have not been fabricated

## Slide 19: Be Isolation Window
0.008” Be window
Cooled by self contained water chiller to 10C
Integrated collimator Densimet

## Slide 20: Cryo-System
ESR 15K He for cooling
Must return ~25K
PID Control heater
  40K
Return mixed with bypass.
Bypass valve on PID
Alarms/interlocks on TS-5 (a/b) and TS-6 (a/b)
Similar controls to Cryotarget
Status 20% Complete

## Slide 21: Shipping
T2 is flammable and radioactive
  Applicable regulations come from both DOT and NRC/DOE
Cell is supported in custom carrier rack
Rack is installed in MTV
  Miscellaneous Tritium Vessel
MTV is shipped in BTSP
  BTSP = Bulk Tritium Shipping Package
Approval to use
  BTSP/MTV granted in July 2017
JLAB T2 was to be first to use MTV

## Slide 22: Carrier Rack/Cell Assembly
Completed Rack Assembly used for testing/training
Specific Procedures to use Rack/MTV are under development

## Slide 23: Controls and Safety Systems
Control System Similar to Hall A Cryo-Target Controls
Numerous FSDs/Alarms
Tritium Monitors
  On loan from SRS/SRTE
  Handheld and Remote
  Remote air suction tubing has been installed
  T2 monitors are being integrated into controls
Monitors will issue FSD and activate exhaust system
Software is under development
Bio Screening program has been initiated
  Limited number of participants

## Slide 25: Exhaust System Activation

## Slide 26: Vacuum Fault System
FSD issued on all faults
Getter operates on cold cathode failure
Pressure switch closes main turbo valves
Flow switch on water cooling for Be window
All faults trip alarms on UI
Pressure events trip heater power.

## Slide 27: Administration
Reviews
  ERR Preliminary Design June 2010: Produced more than 50 action items
  ERR Design Review September 2015: Positive comments design considered acceptable
  ERR Experimental Review March 2016: Addressed issues from September adequately
  ERR Final Review for implementation/walkthrough: Needs to happen near the start of the experiment
  SRS Safety Basis for filling completed June 2016
  SRS Safety basis for shipping March 2017
  Facilities Review and Coordination April 2016
DOE MTV/BTSP Approval July 2017
DOE NNSA Approval for T2 transfer December 2016
DOE NEPA Approval August 2016

## Slide 28: Summary
All reviews are nearly complete
Control and Cryo systems are under development
  Moved to lower priority with new schedule
Exhaust system is nearly complete
Shipping issues have been resolved
Safety systems are nearly complete currently testing
Cells are complete
  Working on Spares now
Most administrative tasks are complete
Target should be ready for Fall 2017 run

## Slide 29: Materials-2
Al 7075 is unlisted
  Design basis
    for tension
    = 80% of 24 ksi for shear
    = 150% OF 24 ksi bending
Other wetted materials are SST
  304/304L
  316/316L
  ER316L Filler for welds

## Slide 30: Load Conditions
At room temp
  P = 200 psi
At 40K Beam Off
  Pressure = ~30 psi
  Max Temperature = 40K
Beam On
  Pressure = 36 psi (avg temp of T2 = 53K)
  Max Temperature = ~125K
Cyclic loads
  Cool down/warm up operating cycles = 20
  17800 beam trips (cycles between Beam On and Off)
    150 days, 33% duty factor, 15 trips/hr

## Slide 31: Analysis
Applicable Code ASME B31.3 (2014)
  Section 304.7.2 because of odd geometry
  Used both Hand Calculations and FEA
  Analysis conforming to ASME BPVC VIII D2 with load factors from B31.3 (i.e. 3 instead of 2.4 on P)
Used cyclic screening analysis from D2
  Depth of loads do not require a fatigue analysis
    175 psi pressure cycle (it is closer to 10 psi)
    Considers temperature cycle from 40K to 125K
Design pressure 675 psi
No source of overpressure
Calculations: TGT-CALC-103-002, 7, 8, 12, 13, 14, 15, 17.

## Slide 32: Thermo-Mechanical Model-1
Full temperature load
  Beam on at 20 µA 2x2 mm raster
  Pressure load 400 psi internal (more than 10x)
  Cooling using 40K heat sink
Using an elastic-plastic model
  Model solves and stresses are still below allowable even for over conservative case
  Local plastic failure requirements met
  Analysis not required because of screening analysis

## Slide 33: Thermo-Mechanical Model-2

## Slide 34: Thermo-Mechanical Model-3

## Slide 35: Mech Therm Model of Raster Off

## Slide 36: Raster Off Time Dependence

## Slide 37: Density Model
T2 properties derived from H2
  Viscosity, Thermal Conductivity, Heat Capacity, etc.
  Assumed a Real Gas model
  Buoyancy, convection on wall included
Assumed fixed 2.8W from 20 µA and 2x2 mm raster (11 mW/mm linear power density)
  Did not correct heat load for density
Averaged 20% reduction in density along beam path

## Slide 38: Vacuum System
Scattering chamber (standard Hall A)
  1900 liters
  Thin sections for recoil particles (0.014” aluminum)
Two 800 l/s turbos backed by Leybold D60 Mech pump
NEG Pump with backing turbo and mech pumps
Vacuum exhaust part of Tritium Exhaust System and is continuously purged with N2 (1 cfm)
Isolated from upstream beamline vacuum (Be window)
Remote RGA may help diagnose leaks. Serve as leak detector.

## Slide 40: Materials
Main Body and Entrance
  Aluminum 7075-T651 ASTM B209
  Extensive use of this allow for 15 years
  Strong, ductile, hard, non weldable
Seals are Al 1100
Valve assembly
  SST 304/304L Fitting
  Swagelok valve all metal bellows sealed (316L)
  Butt welded ER316L (100% VT in process and RT)

## Slide 41: Beam Heating
Max beam current min raster
3W in Entrance
3.3 W in Exit
on exit on entrance

## Slide 42: Summary of Test Results
Multiple hydrotests on components and assemblies
Entrance: Minimum burst above 2900 psi
Main body: Minimum burst above 3400 psi (0.014” section)
Assembly with covers
  Leaked above 4000 psi (seal was damaged)
  Failed above 5500 psi

## Slide 43: Raster Off
Initial conditions:
  Beam on full raster
  20 microA
Exit window is worst case
Beam spot no raster
  0.150 mm diameter
  Square profile
For high energy beam this spot size is very conservative
Tune shall be checked at each energy
Long term operations at these conditions are forbidden
Typical FSD for raster failure
  < 10 ms
Red curve is upstream section temperature

## Slide 44: General Requirements
Collect T2 from any release inside Hall A and exhaust in a controlled fashion
Exhaust point shall be 20 m above grade at site boundary
Must serve as part of the smoke removal system (at least 1/3 of the 36000 cfm required)
Must have at least two modes to service hut and to exhaust from Hall A. (500 and 12000 cfm)
Must stack vacuum pump exhaust
  Scattering chamber, dump line, getter system
Makeup air comes from new louvered door at bottom of ramp.
  Prevents overpressure on ramp door.
  Test required for louver system
  Air from outside from smoke removal system on ramp with damper removed

## Slide 45: Exhaust Summary
Exhausts secondary (vacuum chamber) and tertiary (Hall A) containment to 20 m stack
Two speed
  ~500 cfm
  12000 cfm
Exhaust system activated
  Vacuum switch failure (interlock)
  Truck ramp lower door (interlock)
  Manual activation (Hall and Counting House) (manual)
  Low speed activated manually for hut (manual)
  T2 monitor (interlock)

## Slide 46: Status
Successful review on 15 September 2015
  Report available on wiki
  Pressure system calculations reviewed by 2/1/2016
  Final ladder assembly fixed by 3/1/2016
  ERR 3/16/2016
Design of platform complete
Construction of handling hut complete
Cell design and prototyping/testing complete
  Fabrication underway for actual cells
Modifications to Qweak target underway
  Cryo sytem/piping ~50% complete
Material testing is underway at SRS/SRTE/SRNL
Working through unresolved issues regarding shipping
NEPA analysis is complete should be presented to our site office by 2/1/2016

## Slide 47: Schedule
Hut and Platform
  Fabrication complete by 3/1/2016
  Test install June 2016
T2 Exhaust system
  Internal vent install April 2016
  External May and June 2016
Target
  Complete modifications to Qweak 3/1/2016
  Assembly complete 8/1/2016
  T2 cell install 10/1/2016
Interlocks
  Checkout complete 8/30/2016
ERR 3/16/2016
Checkout review 9/1/2016
Dates are estimates