## Slide 1: Sensitive Ion Pump Current Monitoring using an In-House Built Ion Pump Power Supply
ABSTRACT
UHV supplies monitor ion pump current to 0.1nA, corresponding to ~10-11 Torr
Supplies used on CEBAF injector baked beamline monitor vacuum activity associated with
  Onset of field emission during gun high voltage processing
  Viewer discharge events when running from unanodized photocathode with beam halo
Provide additional laboratory diagnostic for pressure in UHV systems
J. Hansknecht, P. Adderley, A. Comer, M. L. Stutzman
Chart: Ion Pump Current vs. Extractor Gauge Pressure
| 7.7000000000000128E-12 | 2.3000000000000034E-11 | 4.1000000000000046E-11 | 5.9000000000000119E-11 | 9.4000000000000218E-11 | 2.1000000000000039E-10 | 3.1000000000000058E-10 | 5.6000000000000075E-10 | 1.2000000000000019E-9 | 1.900000000000003E-9 | 4.9000000000000083E-9 | 7.3000000000000099E-9
H2 bleed in | 0.24 | 1.1 | 2.4 | 4.1 | 8.1 | 29.5 | 47.7 | 75 | 120 | 350 | 1300 | 2170
Nitrogen | 0.17 | 0.26 | 0.4 | 0.5 | 0.7 | 1 | 1.45 | 1.5 | 1.6 | 1.8 | 2.1 | 3 | 6 | 7.5 | 12.1 | 19 | 22.8 | 35.2 | 52 | 91 | 110 | 145 | 155 | 310 | 530 | 750 | 1240 | 1300 | 2250 | 4100 | 4300 | 4600
Hydrogen | 0.25 | 0.28 | 0.65 | 0.94 | 1 | 1.11 | 1.16 | 1.47 | 1.53 | 1.64 | 1.96 | 2.27 | 2.74 | 3.1 | 3.15 | 4.04 | 4.26 | 4.5 | 5.44 | 6.5 | 7.07 | 7.56 | 8.52 | 9.07
H2 again day 2 | 0.26 | 3 | 3.82 | 4.6 | 6.16 | 8 | 8.4 | 10.7 | 12.2 | 17.4 | 21.1 | 42.2 | 64 | 80 | 107 | 350 | 880 | 1250 | 1500 | 2620 | 3150 | 3580
marcy nitrogen 7/28 | 0.23 | 4 | 5.95 | 3.5 | 2.2 | 0.76 | 0.47 | 0.99 | 1.33 | 1.88 | 2.03 | 2.61 | 4.42 | 4.91 | 6.21 | 6.72 | 12.9 | 14.5 | 17.7 | 20 | 21.7 | 29.1 | 61.1 | 75.4 | 84.4 | 124 | 150 | 172 | 191 | 300 | 667 | 1020 | 1600
2003 Hydrogen data | 10 | 20 | 100 | 150 | 300 | 400 | 300
1-28-04 ext2 | 15 | 30 | 200 | 350 | 400 | 450
ext2 2-5-04 | 15 | 20 | 30 | 1000 | 3000 | 4000 | 5000 | 5000
ext 11-20-03 | 100000 | 1000 | 100
hydrogen, BeCu, 4 kV | 19.8 | 425 | 486 | 446 | 696 | 910 | 1280 | 3000 | 6980 | 8740 | 16100 | 18700 | 24400 | 39600 | 97000
Manufacturer Data | 1200 | 6000 | 50000 | 600000 | 5000000
~10-10 Torr
Full Scale
Discharge event in beamline
Current (10-10A)
UHV supply current during RGA scan. Sawtooth period corresponds to RGA scan frequency.
RGA on and scanning
RGA off
Notice: Authored by Jefferson Science Associates, LLC under U.S. DOE Contract No. DE-AC05-06OR23177.
Chart: Pressure and Current vs. Voltage
| 7000 | 6355 | 5710 | 5065 | 4420 | 3775 | 3130 | 2485 | 1840 | 1195 | 550
Pressure (Torr) | 1.78e-10 | 1.66e-10 | 1.54e-10 | 1.36e-10 | 1.2e-10 | 1.1e-10 | 1.02e-10 | 9.4e-11 | 9.4e-11 | 1.1e-10 | 2.56e-10
Ion pump current | 76.7 | 67.1 | 53.4 | 45.5 | 36.7 | 29.5 | 23.9 | 18.2 | 14.3 | 11 | 8.69
Voltage dependence
Lowering voltage past threshold results in reduction in pumping, evidenced by higher extractor gauge pressure
Onset of field emission
during HV processing