TJNAF/Hall-B

PrimEx Note 25

HYCAL Data Acquisition Live Time Analysis

D. Pomeroy

University of Massachusetts Amherst

July 22, 2004

1.  Introduction

The data acquisition system (DAQ) constructed for the Primakoff Experiment (TJNAF E00-014) currently consists of three fastbus crates with modules which read the output of 31 ADCs and 1 TDC.  A fourth crate to control the readout of tagger electronics will be added upon installation into Hall B.  The read out of these modules is controlled by the Trigger Supervisor. In any data acquisition system there is a certain portion of time in which the system is unable to take inputs as it is processing previous data.  The percentage of time in which the system is able to take inputs is known as the live time.  As the event rate is increased the live time decreases and therefore limits the rate at which data can be recorded.  The following studies explore three different aspects of the DAQ.  The first shows a full read out all available channels, the second is a read out of approximately 100 channels, and the third is a read out of a varied number of channels.  The tests were performed to investigate the limits of the data rate of the DAQ, and show a maximum data rate under run time conditions is ~4 kHz and will have a live time of ~85%.

2.  Two Fast-Bus Read Out

The first situation studied was the read out of just two fastbus crates (as the third was in repair).  It is possible to set the threshold on ADCs in order to ensure that signals read out are of significant size to be considered real data.  This threshold is known as the sparsification threshold and for this portion of the tests was turned off.  This created a situation in which for each event every channel was read out, or approximately 2/3 of the total channels (since only 2 of the 3 crates were working).  The trigger rate was then varied from 0.1 to approximately 6.0 kHz.  For each run the integrated values for the live time, event rate, and data rate were recorded (see Table 1).  In Figure 1 the live time vs. data rate was plotted.  As the plot indicates, data rates higher than ~9000 kb/second experience a sharp drop in live time.  In addition, a plot of the data rate vs. trigger rate can be used to determine what trigger rate corresponded to the maximum data rate of 9000 kb/second (see Figure 2).  In this case the data rate increased linearly with the trigger rate until the trigger rate reached approximately 2 kHz.  This shows that for a full read out of two fastbus crates the maximum trigger rate is ~2 kHz which produces a data rate of 9000 kb/second without significant loss of live time.  This result is important for tests of such as the Light Monitoring System or pedestal runs, as it shows that the trigger rates should be kept under 2 kHz.

3.  ~100 Channel Read Out

The second study was to try to pick a threshold so that only about 100 channels were being read out for every event.  This was chosen because a two cluster event should contain approximately 100 channels, and this is exactly what one would expect from a 0 event.  To do this, it was assumed that each event would contain about 6 bytes of data, so the sparsification threshold was set such that the event size was ~600 bytes, or a threshold of 640.  The trigger rate was varied from approximately 2.5 kHz to 20 kHz and the integrated values for live time, event rate, and data rate were recorded (see Table 2).  The live time was plotted as a function of data rate and shows a cutoff ~2300 kb/second (see Figure 3).  When live time was plotted against trigger rate, it showed the cutoff correlating with a trigger rate of about 4 kHz (see Figure 4).  This means that with the current triggering scheme it would be possible to accept triggers at a rate of up to 4 kHz.

4.  Sparsification Threshold Variation

In order to assure the trigger rate is kept far from the cut off the experiment will run at a rate of approximately 3.3 kHz.  In the final situation studied the sparsification threshold was varied while holding the trigger rate at 3.3 kHz.  Because the gain in each channel is different the higher the sparsification threshold is set the fewer channels will be read out for each event.  This will cause a decrease in the data rate.  To do this the sparsification thresholds were varied from 300 to 700.  This time the live time, data rate, event rate, and approximate event size were recorded (see Table 3).  From this again live time can be plotted as a function of data rate to show at what data rates significant loss in live time occurs (see figure 5).  Here the cutoff appears to be at ~5000 kb/second.  This shows that for an input trigger rate of 3.3 kHz the number of channels read out in an event can be increased to read out a maximum of 5000 kb/second.

5.  Conclusion

The data acquisition system for the Primakoff Experiment consists of a variety of electronics which have fundamental limits on the speed at which they can read data.  In order to test these limits three different assessments were performed.  First the live time was measured as the trigger rate was varied with no sparsification thresholds set for two fastbus crates.  This showed a maximum data rate of 9000 kb/second at a maximum trigger rate of 2 kHz.  Second the sparsification threshold was set to imitate an event size of approximately 100 channels and it was shown that the maximum data rate was 2300 kb/second at a trigger rate of 4 kHz.  Finally the sparsification threshold was varied while the trigger rate was set at 3.3 kHz producing a maximum data rate of 5000 kb/second.

Tables

| Trigger Rate (Hz) | Live Time (%) | Event Rate | Data Rate |
| --- | --- | --- | --- |
| 5882 | 16.9 | 1888 | 9812 |
| 5000 | 19.7 | 1901 | 9817 |
| 4545 | 21.8 | 1898 | 9803 |
| 4000 | 24.1 | 1873 | 9670 |
| 3448 | 28.1 | 1907 | 9846 |
| 3030 | 32.4 | 1893 | 9777 |
| 2703 | 36.0 | 1842 | 9514 |
| 2632 | 43.3 | 1860 | 9605 |
| 2500 | 48.8 | 1886 | 9757 |
| 2381 | 52.4 | 1855 | 9580 |
| 2326 | 56.5 | 1854 | 9574 |
| 2222 | 61.3 | 1845 | 9527 |
| 2083 | 74.4 | 1895 | 9785 |
| 2000 | 85.4 | 1911 | 9868 |
| 1887 | 89.0 | 1832 | 2045 |
| 1786 | 90.7 | 1761 | 9095 |
| 1695 | 91.2 | 1668 | 8612 |
| 1587 | 91.9 | 1557 | 8038 |
| 1493 | 93.2 | 1185 | 6117 |
| 1099 | 95.2 | 1086 | 5610 |
| 1000 | 94.5 | 988 | 5103 |
| 901 | 95.3 | 897 | 4632 |
| 800 | 95.0 | 789 | 4073 |
| 699 | 95.5 | 690 | 3563 |
| 599 | 97.2 | 588 | 3038 |
| 500 | 97.2 | 495 | 2555 |
| 400 | 97.6 | 398 | 2059 |
| 299 | 98.5 | 298 | 1540 |
| 200 | 99.0 | 198 | 1023 |
| 100 | 99.5 | 100 | 516 |
| Table 1 Trigger Rates were varied between 0.1 and 6 kHz. Live time, event and data rates were then recorded. |

| Trigger Rate | Live Time | Event Rate | Data Rate |
| --- | --- | --- | --- |
| 20000 | 12.6 | 4202 | 2413 |
| 11111 | 20.2 | 4188 | 2403 |
| 8333 | 29.2 | 4126 | 2366 |
| 7143 | 32.2 | 4213 | 2467 |
| 6250 | 35.0 | 4110 | 2403 |
| 5556 | 41.9 | 4170 | 2437 |
| 5000 | 49.1 | 4152 | 2430 |
| 4762 | 61.8 | 4202 | 2458 |
| 4545 | 90.0 | 4133 | 2415 |
| 3448 | 96.1 | 3305 | 1931 |
| 2564 | 96.5 | 2467 | 1445 |
| Table 2 With the sparsification threshold set to 640 the trigger rate was varied from ~2.5 kHz to 20 kHz. The integrated values for live time, event rate and data rate were then recorded. |

| Threshold | Live Time | Event Rate | Data Rate | Event Size |
| --- | --- | --- | --- | --- |
| 300 | 14.4 | 898 | 6050 | 7000 |
| 400 | 24.6 | 1195 | 5800 | 3000 |
| 425 | 30.9 | 1304 | 5618 | 4400 |
| 450 | 37.7 | 1531 | 5645 | 3500 |
| 475 | 51.9 | 1942 | 5945 | 3300 |
| 485 | 52.1 | 1929 | 5461 | 2900 |
| 495 | 66.3 | 2362 | 6065 | 2650 |
| 500 | 74.0 | 2656 | 6512 | 2500 |
| 525 | 92.1 | 3091 | 5790 | 1900 |
| 550 | 92.1 | 3250 | 4740 | 1500 |
| 575 | 95.8 | 3266 | 3890 | 1225 |
| 600 | 97.5 | 3278 | 3078 | 900 |
| 700 | 95.3 | 3231 | 1053 | 330 |
| Table 3 With the trigger rate set to ~3.3 kHz the sparsification threshold was varied between 300 and 700. The integrated values for live time, event rate, data rate and event size were recorded. |

# Text boxes

Figure 1

Live Time vs Data Rate with sparsification thresholds turned off.  From this plot it is clear that there is significant loss in live time for data rates higher than 9000 kb/second.

Figure 2

Data Rate vs. Trigger Rate with sparsification thresholds turned off.  The data rate increases linearly with the trigger rate until it plateaus at 9000 kb/second which shows a maximum trigger rate of ~2 kHz.

Figure 5

Live Time vs. Data Rate with a trigger rate of ~3.3 kHz.  The sharp cut off of the live time shows a maximum possible data rate of 5000 kb/second.

Figure 4

Live Time vs. Trigger Rate with sparsification threshold set to 640.  This shows that for a 100 channel read out the maximum trigger rate will be approximately 4 kHz.

Figure 3

Live Time vs. Data Rate for a sparsification threshold of 640.  This plot shows a clear maximum data rate for approximately 100 channels at ~2300 kb/second.