| Jefferson Lab Alignment Group Data Transmittal |
| --- |
| TO: D. Kashy | DATE: Jul 14 2004 |
| FROM: J. Dahlberg | Checked: | # : B936 |

DETAILS:

Below are the results from the Jan 28, 2004 inspection performed on the HyCal detector located in the Test lab.  Note that the inside edges of the crystals near centerline at the downstream end were difficult to reach and the beam left and top edge readings are questionable.  Two sets of data are listed, the first with a coordinate system based on the circular aperture at beam centerline, and the second relative to average planes defined off the inside edge of the crystals.  This gives a larger lever arm to set pitch and yaw.  The measurements on the crystals at centerline were not used to define the second coordinate system because of the short lever arm and the difficulty to reach. The roll for both coordinate systems was set using an average of the four inside planes defined off the crystals.  A +X is to the beam left, and +Y is above.  The measurements on the outside most crystals are numbered as labeled on the detector.

DOWNSTREAM END

LOCATION		 	    X		    Y			    X		    Y

CL APERTURE	 		    0.00		    0.00			    0.17		    0.25

INSIDE EDGES OF CRYSTALS

BL 1				 351.33					 351.68

BL 8				 351.41					 351.72

BL 17				 351.97					 352.25

BL 26				 351.59					 351.84

BL 34				 351.46					 351.67

BR 1				-354.30					-353.94

BR 8				-353.65					-353.34

BR 17				-352.80					-352.52

BR 26				-352.64					-352.39

BR 34				-352.73					-352.51

TOP 1						 352.45					 352.85

TOP 8						 352.48					 352.85

TOP 17						 352.50					 352.85

TOP 26						 352.78					 353.11

TOP 34						 352.26					 352.55

BOT 1						-352.80					-352.38

BOT 8						-352.83					-352.44

BOT 17						-353.03					-352.67

BOT 26						-352.96					-352.64

BOT 34						-353.02					-352.73

INSIDE EDGE OF CRYSTALS AT CENTER APERTURE

BR BOT						  -20.83					  -20.58

BL BOT						  -20.94					  -20.68

BL SIDE BOT			   16.41					   16.58

BL SIDE TOP			   19.98					   20.14

BL TOP 						   19.41					   19.67

BR TOP						   20.97					   21.23

BR SIDE TOP			  -20.73					  -20.57

BR SIDE BOT			  -20.92					  -20.75

| Jefferson Lab Alignment Group Data Transmittal Continued | Page 2 |  | of 2 |  |
| --- | --- | --- | --- | --- |
| Date : | July 14, 2004 | Transmittal # : | B936 |

UPSTREAM END

LOCATION		 	    X		   Y			    X		  Y

CL APERTURE			    0.00		   0.00			    0.07		   0.17

INSIDE EDGES OF CRYSTALS

TOP 1						352.48					352.52

TOP 86						353.10					353.21

TOP 136						353.15					353.23

TOP 142						352.68					352.83

TOP 173						352.52					352.66

TOP 213						353.06					353.15

TOP 651						353.14					353.20

TOP 728						351.80					351.97

BR 155				-352.51					-352.47

BR 158				-353.68					-353.75

BR 202				-353.06					-353.06

BR 227				-352.25					-352.18

BR 267				-352.80					-352.78

BR 532				-353.46					-353.51

BR 875				-353.28					-353.30

BR 936				-352.42					-352.36

BOT 321						-353.13					-353.02

BOT 331						-352.87					-352.77

BOT 419						-353.47					-353.31

BOT 422						-353.16					-353.03

BOT 649						-352.92					-352.87

BOT 680						-353.62					-353.45

BOT 845						-352.86					-352.78

BOT 914						-352.97					-352.91

BL 111				 353.58				353.53

BL 151				 353.29				353.35

BL 298				 354.01				354.03

BL 569				 353.40				353.39

BL 595				 353.66				353.58

BL 622				 353.54				353.51

BL 854				 353.73				353.73

BL 990				 353.83				353.87

INSIDE EDGE OF CRYSTALS AT CENTER APERTURE

BL TOP						   20.67					   20.84

BR TOP						   20.82					   20.98

BR SIDE TOP			  -20.51				-20.45

BR SIDE BOT			  -20.57				-20.50

BR BOT						  -20.91					  -20.75

BL BOT						  -20.94					  -20.77

BL SIDE BOT			   21.07				 21.14

BL SIDE TOP			   21.11				 21.17