| Jefferson Lab Alignment Group Data Transmittal |
| --- |
| TO: Mark Ito | DATE: Sep 10 2004 |
| FROM: Chris Gould | Checked: | # : B944 |

DETAILS:

Below are the results of the Primex Target Calibration carried out September 1, 2004.   A right hand coordinate system was established with the Z axis running through the center of the upstream flange.  Positive X is to the beam left, positive Y is up.  Values are in millimeters.  Attached is a sketch to provide the naming convention.  All features shown were measured in the initial “Start” position to show geometry.  All additional measurements were made to show relative movements from the “Start” position.   The As-Set position of the target relative to beam line is X = -0.63, Y = 0.70, Z =-14754.7 and should be used to calculate the offsets of the following measurements.

## Position	Description		      X		    Y

Start               MH 1                              -4.43          -10.59

MH 2                              -4.33          -28.88

MH 3                            -34.63          -10.82

MH 4                            -34.47          -28.96

MH 5                            -64.73          -10.86

MH 6                            -64.69          -29.04

MH 7                            -94.88          -10.89

MH 8                            -95.23          -28.98

MH 9                          -125.13          -10.95

MH 10                        -125.05          -29.13

TG 1                            -19.46           -39.02

TG 2                            -49.63           -39.08

TG 3                            -79.83           -39.13

TG 4                          -110.02           -39.15

TG 5                          -149.07           -22.48

20mm Left      MH 1                            15.48           -10.56

MH 2                            15.62           -28.83

MH 3                           -14.72           -10.84

MH 4                           -14.56           -28.96

MH 5                           -44.81           -10.88

MH 6                           -44.77           -29.05

### MH 7                           -75.00           -10.92

MH 8                           -75.15           -29.07

MH 9                         -105.21           -10.90

MH 10                       -105.13           -29.12

TG 5                          -129.16           -22.51

| Jefferson Lab Alignment Group Data Transmittal Continued | Page | 2 | of | 5 |
| --- | --- | --- | --- | --- |
| Date : | September 10, 2004 | Transmittal # : | B944 |

## Position	Description		       X		     Y

Start Rpt         MH 1                            -4.46          -10.61

MH 2                            -4.33          -28.88

MH 3                          -34.65          -10.85

MH 4                          -34.47          -29.00

MH 5                          -64.73          -10.89

MH 6                          -64.69          -29.09

MH 7                          -94.92          -10.90

MH 8                          -95.11          -29.09

MH 9                        -125.16          -10.89

MH 10                      -125.09          -29.13

TG 5                         -149.10          -22.54

178mm Left	MH 1                           172.69         -10.57

MH 2                           172.78         -28.90

MH 3                           142.48         -10.81

MH 4                           142.67         -28.94

MH 5                           112.36         -10.85

MH 6                           112.46         -29.02

MH 7                             82.22         -10.86

MH 8                             82.07         -29.03

MH 9                             52.00         -10.83

### MH 10                           52.06         -29.07

TG 5                              27.99         -22.50

90mm Left	MH 1                              85.39        -10.58

MH 2                              85.45        -28.91

MH 3                              55.19        -10.80

MH 4                              55.34        -28.95

MH 5                              25.06        -10.84

MH 6                              25.11        -29.04

MH 7                               -5.10        -10.90

MH 8                               -5.28        -29.04

MH 9                             -35.36        -10.90

MH 10                           -35.25        -29.09

TG 5                              -59.29        -22.45

| Jefferson Lab Alignment Group Data Transmittal Continued | Page | 3 | of | 5 |
| --- | --- | --- | --- | --- |
| Date : | September 10, 2004 | Transmittal # : | B944 |

## Position	Description		       X		     Y

12mm Up        MH 1                             -4.44             1.18

MH 2                             -4.37          -17.11

MH 3                           -34.65             0.96

MH 4                           -34.48          -17.18

MH 5                           -64.80             0.94

MH 6                           -64.71          -17.28

MH 7                           -94.91             0.87

MH 8                           -95.11          -17.30

MH 9                         -125.13             0.88

MH 10                       -125.05          -17.37

TG 5                          -149.09          -10.73

12mm Up Rpt MH 1                             -4.44             1.43

MH 2                             -4.35          -16.84

MH 3                           -34.66             1.21

MH 4                           -34.49          -16.91

MH 5                           -64.81             1.15

MH 6                           -64.69          -17.03

MH 7                           -94.92             1.14

MH 8                           -95.08          -17.06

MH 9                         -125.17             1.09

MH 10                       -125.09          -17.13

TG 5                          -149.06         -10.42

Start Rpt         MH 1                              -4.43         -10.52

MH 2                              -4.32         -28.80

MH 3                            -34.66         -10.77

MH 4     		    -34.47         -28.91

MH 5   		    -64.74         -10.82

MH 6    		    -64.72         -29.01

MH 7      		    -94.92         -10.87

MH 8			    -95.11         -28.99

MH 9                          -125.13         -10.86

MH 10                        -125.09         -29.10

TG5                            -149.06         -22.41

| Jefferson Lab Alignment Group Data Transmittal Continued | Page | 4 | of | 5 |
| --- | --- | --- | --- | --- |
| Date : | September 10, 2004 | Transmittal # : | B944 |

## Position	Description		       X		     Y

25mm Up        MH 1                              -4.45           14.47

MH 2                              -4.38            -3.84

MH 3                            -34.66           14.22

MH 4                            -34.50            -3.89

MH 5                            -64.79           14.18

MH 6                            -64.73            -3.99

MH 7                            -94.94           14.18

MH 8                            -95.08            -4.02

MH 9 			  -125.15           14.15

MH 10                        -125.09            -4.10

TG 5                           -149.08             2.54

90mm Left      MH 1                             85.46           -10.62

Rpt            MH 2                             85.50           -28.94

MH 3                             55.23           -10.84

MH 4                             55.37           -29.02

MH 5                             25.11           -10.94

MH 6                             25.17           -29.10

MH 7                              -5.09           -10.95

MH 8                              -5.22           -29.10

MH 9                            -35.28           -10.92

MH 10                          -35.19           -29.17

TG 5                             -59.19           -22.49

178mm Left X MH 1                            172.70            14.62

25mm Up      MH 2                            172.79             -3.66

MH 3                            142.47            14.39

MH 4                            142.66             -3.75

MH 5                            112.35            14.34

MH 6                            112.45             -3.85

MH 7                              82.23            14.34

MH 8                              82.08             -3.88

MH 9                              51.97            14.28

MH 10                            52.05             -3.94

TG 5                             28.06               2.74

| Jefferson Lab Alignment Group Data Transmittal Continued | Page | 5 | of | 5 |
| --- | --- | --- | --- | --- |
| Date : | September 10, 2004 | Transmittal # : | B944 |