## Slide 1: SRF Science and Technology
Grigory Eremeev
20 June 2019

## Slide 2: Jefferson Lab Overview
Core Competencies
  Accelerator Science and Technology
  Large Scale User Facilities/Advanced Instrumentation
  Nuclear Physics
Mission Unique Facilities
  Continuous Electron Beam Accelerator Facility

## Slide 3
CEBAF @ 12 GeV
New Hall
Add arc
Enhanced capabilities
in existing Halls
Add 5 cryomodules
Add 5 cryomodules
20 cryomodules
20 cryomodules
Upgrade arc magnets
and supplies
CHL upgrade
Constructed 1987-1993: ~ 1 km long
First large high-power CW recirculating e-linac based on SRF technology
Recent energy upgrade to 12 GeV
  10 new high performance cryomodules
  Double the capacity of the Central Helium Liquefier
  Upgrade magnets and power supplies for recirculation arcs
  Upgrade Extraction, Instrumentation and Diagnostics, and Safety Systems
  Add new beamlines for Arc 10 and Hall D
  Add new experimental Hall D and upgrade existing Halls

## Slide 4
SRF technology photo gallery

## Slide 5
SRF technology photo gallery

## Slide 6: Cavity fundamentals
Conducting cavities can be filled with electromagnetic energy. This energy is then transferred to charged particles flying along the cavity axis.
Lowest frequency eigen-mode has a longitudinal electric field  use this mode for acceleration
Typical RF frequency: 0.4 GHz to 4 GHz
Higher-Order (higher frequency) modes: unwanted  need damping

## Slide 7
SRF globally
CEA
Orsay
KEK
LANL
INFN Legnaro
INFN Genoa
INFN Milan
TU Darmstadt
Peking
University
Australian
National
University
JLab
CERN
ANL
FNAL
CEA Saclay
DESY/FLASH/XFEL
CESR
TRIUMF
MSU/FRIB
IHEP
ESS
SLAC

## Slide 8: SRF  Superconducting Radio Frequency
SRF science is the study of superconducting RF properties of particular materials as well as designs of accelerator systems to provide understanding of performance limitations in accelerator systems.

## Slide 9: SRF  Superconducting Radio Frequency
SRF technology exploits superconducting materials to provide high performance, energy efficient accelerator systems.
1972: First superconducting accelerator based on L-band 1300 MHz niobium cavities @ Stanford High Energy Physics Lab
CEBAF was the first large-scale application of SRF in the US – constructed 1987-1993: ~ 1 km long
XFEL accelerator in Europe is built 2009-2015: ~ 2 km long
LCLS-II @ Stanford Linear Collider Complex currently under construction 2016-2020: ~ 1 km long, 35 cryomodules
Proposed ILC accelerator 20??-20??: ~ 30 km long, ~ 1000 cryomodules
Essentially all new large accelerators being planned in the world aim to exploit this technology.

## Slide 10
THE CAVITY
Cavity
cavity
Niobium – best superconducting properties among all pure metals:
    Tc ~ 9.25 K;
    Hc ~ 2000 Oe;
    Rbcs ~ .00001 mΩ at 2 K
Nb : Rs ~ .00001 mΩ
Cu: Rs ~ 10 mΩ
↓
Qnb ~ 1011 up to
Eacc = 5·107 Volts per meter
Niobium, the material of choice

## Slide 11
Ideal?
9
10
8
10
11
10
10
10
0
20
40
60
Q0
Eacc, MV/m
Real
Real
Real

## Slide 12
High-purity Nb, sheet inspection, CBP, etc
Improved cavity shape
Electron multiplication
Thermal breakdown
Electron field emission
Quench
High-pressure rinsing, cleanroom procedures
Electropolishing,
low temperature baking, …
Limitation:
Solution:
Limitation:
Solution:
1975
3 MV/m
1995
25 MV/m
1990
15 MV/m
2010
50 MV/m
1985
6 MV/m
2000
40 MV/m
High field Q-reduction
Further improved cavity shapes
“Physics of dirt”, Pauli 1933

## Slide 13: Residual resistance
RF surface resistance of SRF cavities level off to a constant level at low temperatures, which is not predicted by BCS theory of superconductivity

## Slide 14: Residual resistance
Earth magnetic field = 0.5 Gauss => Rs ~ 100 nOhm => Q0 ~ 108
Nb2O5 oxide is a “good” dielectric, but NbO2 and NbO are poor conductors/semiconductors.
Hydrogen is readily absorbed by Nb at room temperature and has high diffusion rate even at low temperature, which leads to formation of lossy NbH clusters below 100 K.
Lattice damage

## Slide 15: Residual resistance
Etch the damage layer and contamination with chemistry
Pushing magnetic fluxiods from high current areas with thermal gradients
High quality factors and low residual resistances < 0.5 nOhm has been achieved.

## Slide 16
High-purity Nb, sheet inspection, CBP, etc
Improved cavity shape
Electron multiplication
Thermal breakdown
Electron field emission
Quench
High-pressure rinsing, cleanroom procedures
Electropolishing,
low temperature baking, …
Limitation:
Solution:
Limitation:
Solution:
1975
3 MV/m
1995
25 MV/m
1990
15 MV/m
2010
50 MV/m
1985
6 MV/m
2000
40 MV/m
High field Q-reduction
High field Q-reduction
Further improved cavity shapes

## Slide 17: Thermal breakdown
Thermal breakdown occurs when the heat generated on RF surface exceeds the heat conduction capacity from RF surface through cavity wall and Nb-He interface to liquid helium bath
Small size particles on RF surface coupled with poor thermal conductivity of Nb wall at low temperatures can cause thermal breakdown

## Slide 18: Thermal breakdown
Defects on niobium surface is a common occurrence especially close to the electron beam weld regions, where niobium melting (Tm=2700 K) and rapid solidification occur.

## Slide 19
High-purity Nb, sheet inspection, CBP, etc
Improved cavity shape
Electron multiplication
Thermal breakdown
Electron field emission
Quench
High-pressure rinsing, cleanroom procedures
Electropolishing,
low temperature baking, …
Limitation:
Solution:
Limitation:
Solution:
1975
3 MV/m
1995
25 MV/m
1990
15 MV/m
2010
50 MV/m
1985
6 MV/m
2000
40 MV/m
High field Q-reduction
High field Q-reduction
Further improved cavity shapes

## Slide 20: Electron field emission

## Slide 21: Electron field emission

## Slide 22: Electron field emission

## Slide 23: Electron field emission

## Slide 24: Electron field emission
A
Wilhelm Röntgen

## Slide 25: Electron field emission
Φ – material work function
k – effective surface area
E – electric field
β – field enhancement factor
Electron current is not expected from Nb surface below GV/m.
Experiments with single-crystal Nb samples show FE onset higher than 1 GV/m
Fowler-Nordheim

## Slide 26: Electron field emission
A. Dangwal et al., PRST-AB 12, 023501 (2009).
A. Navitski et al., PRST-AB 16, 112001 (2013).

## Slide 27: Electron field emission
Intrinsic FE measurements revealed anisotropic values of 4.02 and 3.8 for (111) and (100) orientations of Nb, respectively.

## Slide 28: Electron field emission
Tip-on-tip model

## Slide 29: Electron field emission

## Slide 30: Electron field emission

## Slide 31
High-purity Nb, sheet inspection, CBP, etc
Improved cavity shape
Electron multiplication
Thermal breakdown
Electron field emission
Quench
High-pressure rinsing, cleanroom procedures
Electropolishing,
low temperature baking, …
Limitation:
Solution:
Limitation:
Solution:
1975
3 MV/m
1995
25 MV/m
1990
15 MV/m
2010
50 MV/m
1985
6 MV/m
2000
40 MV/m
High field Q-reduction
Further improved cavity shapes

## Slide 32: High field quench
Reduction of surface magnetic field by increasing magnetic volume

## Slide 33
Ideal?
9
10
8
10
11
10
10
10
0
20
40
60
Q0
Eacc, MV/m
Real
Real
Real
Ideal?

## Slide 34
Unexpectedly, a reduction in resistance (or inversely, an increase in Q) with field have been recently seen on a number of heat treated cavities.

## Slide 35
High-purity Nb, sheet inspection, CBP, etc
Improved cavity shape
Electron multiplication
Thermal breakdown
Electron field emission
Quench
High-pressure rinsing, cleanroom procedures
Electropolishing,
low temperature baking, …
Limitation:
Solution:
Limitation:
Solution:
1975
3 MV/m
1995
25 MV/m
1990
15 MV/m
2010
50 MV/m
1985
6 MV/m
2000
40 MV/m
High field Q-reduction
High field Q-reduction
Further improved cavity shapes
Higher gradients?
?

## Slide 36: Path forward?
Material | Tc (K) | ρn(μΩcm) | Hc(0) [T] | Hc1(0) [T] | Hc2(0) [T] | (0) [nm]
Nb | 9.2 | 2 | 0.2 | 0.17 | 0.4 | 40
NbN | 16.2 | 70 | 0.23 | 0.02 | 15 | 200
NbTiN | 17.5 | 35 |  | 0.03 |  | 151
Nb3Sn | 18 | 20 | 0.54 | 0.05 | 30 | 85
V3Si | 17 |  |  |  |  | 
Mo3Re | 15 |  | 0.43 | 0.03 | 3.5 | 140
MgB2 | 40 |  | 0.43 | 0.03 | 3.5 | 140
50 nm of Nb3Sn

## Slide 37: Bulk Nb3Sn

## Slide 38: Nb on Cu
40 MV/m +
Bulk Nb

## Slide 39: JLab SRF in Summary
State-of-the-art manufacturing, processing and testing facilities for SRF cavities
  Ongoing R&D for LCLS-II/LCLS-II-HE, JLEIC, SNS PPU, and other projects
  Production of LCLS-II cryomodules
  Consulting with others on future projects
World class R&D projects:
  Cavity design and simulation (high gradient, high current)
  Material development (large-grain Nb, thin films, new materials)
  Surface processing (electropolishing, mechanical polishing, etc)
  Basic SRF physics (understanding RF losses in superconductors)
Reduce cost and improve reliability of SRF accelerators worldwide

## Slide 40: Thank you for listening! I hope you have a great summer at Jefferson Lab!