## Slide 1: Scintillation Detectors
Elton Smith -- Jlab Summer Lecture Series
Introduction
Components
  Scintillator
  Light Guides
  Light Sensors
    Photomultiplier Tubes
    Silicon Photomultipliers
Formalism/Electronics
Application to Particle Identification

## Slide 2: Experiment basics
B field ~ 5/3 T
R = 3m
L = ½ p R = 4.71 m
p = 0.3 B R = 1.5 GeV/c
tp = L/bpc = 15.77 ns
tK = L/bKc = 16.53 ns
DtpK = 0.76 ns
bp = p/√p2+mp2 = 0.9957
bK = p/√p2+mK2 = 0.9496

## Slide 3: Experiment basics
B field ~ 5/3 T
R = 3m
L = ½ p R = 4.71 m
p = 0.3 B R = 1.5 GeV/c
tp = L/bpc = 15.77 ns
tK = L/bKc = 16.53 ns
DtpK = 0.76 ns
bp = p/√p2+mp2 = 0.9957
bK = p/√p2+mK2 = 0.9496
Particle Identification by time-of-flight (TOF) requires
Measurements with accuracies of ~ 0.1 ns

## Slide 4: Measure the Flight Time between two Scintillators
400 cm
100 cm
300 cm
20 cm
Disc
Disc
TDC
Start
Stop
450 ns

## Slide 5: Measure the Flight Time between two Scintillators
400 cm
100 cm
300 cm
20 cm
Disc
Disc
TDC
Start
Stop
Particle Trajectory
450 ns

## Slide 6: Propagation velocities
c = 30 cm/ns
vscint = c/n = 20 cm/ns
veff = 16 cm/ns
vpmt = 0.6 cm/ns
vcable = 20 cm/ns
Dt ~ 0.1 ns
Dx ~ 3 cm

## Slide 7: CLAS detector with FC pulled apart

## Slide 8: Start counter assembly

## Slide 9: Scintillator types
Organic
  Liquid
    Economical
    messy
  Solid
    Fast decay time
    long attenuation length
    Emission spectra
Inorganic (crystals)
  Anthracene
    Unused standard
  NaI, CsI
    Excellent g resolution
    Slow decay time
  Lead Tungstate (PbWO4)
    High density and resolution

## Slide 10: Solid scintillators, examples
EPIC Crystal Co., China
Eljen Technologies
plastic

## Slide 11: Scintillator types
Organic
  Liquid
    Economical
    messy
  Solid
    Fast decay time
    long attenuation length
    Emission spectra
Inorganic (crystals)
  Anthracene
    Unused standard
  NaI, CsI
    Excellent g resolution
    Slow decay time
  Lead Tungstate (PbWO4)
    High density and resolution

## Slide 12: Scintillators: Light Spectrum
Photomultipliers and Accessories
Electron Tubes Ltd (1996)

## Slide 13: Light Collection: Light guides
Goals
  Match (rectangular) scintillator to (circular) pmt
  Optimize light collection for applications
Types
  Plastic
  Air
  None
  “Winston” shapes

## Slide 14: Light guides, examples
Eljen Technologies
Techniques for Nuclear and
Particle Physics, W.R. Leo

## Slide 15: Reflective/Refractive boundaries
acrylic
Scintillator
n = 1.58
PMT glass
n = 1.5

## Slide 16: Reflective/Refractive boundaries
acrylic
Scintillator
n = 1.58
PMT glass
n = 1.5

## Slide 17: Reflective/Refractive boundaries
mirror
Scintillator
n = 1.58
PMT glass
n = 1.5
Air

## Slide 18: Reflective/Refractive boundaries
mirror
Scintillator
n = 1.58
PMT glass
n = 1.5
(reflectance at normal incidence)
Air

## Slide 19: Reflective/Refractive boundaries
Scintillator
n = 1.58
PMT glass
n = 1.5
air

## Slide 20: Reflective/Refractive boundaries
Scintillator
n = 1.58
PMT glass
n = 1.5
air

## Slide 21: Reflective/Refractive boundaries
Scintillator
n = 1.58
PMT glass
n = 1.5
air

## Slide 22: Reflective/Refractive boundaries
acrylic
Scintillator
n = 1.58
PMT glass
n = 1.5
Acceptance of incident rays at fixed angle depends
on position at the exit face of the scintillator

## Slide 23: Reflective/Refractive boundaries
acrylic
Scintillator
n = 1.58
PMT glass
n = 1.5
Acceptance of incident rays at fixed angle depends
on position at the exit face of the scintillator

## Slide 24: Reflective/Refractive boundaries
acrylic
Scintillator
n = 1.58
PMT glass
n = 1.5
Large-angle
ray lost
Acceptance of incident rays at fixed angle depends
on position at the exit face of the scintillator
Rule of thumb: Acceptance is given by the ratio of output/input areas

## Slide 25: Winston Cones - special geometry
Rev. Sci. Instrum. 41 (1970) 413-418

## Slide 26: Photomultiplier tube, sensitive light meter
Photocathode
Electrodes
N Dynodes
Anode
56 AVP pmt
Gain ~ VN ~ 106 - 107

## Slide 27: Photomultiplier tube, sensitive light meter
Photocathode
Electrodes
N Dynodes
Anode
56 AVP pmt
g
e−
Gain ~ VN ~ 106 - 107
Photomultiplier Tubes
Principles and Applications, Photonis (2002)

## Slide 28: Examples from Hamamatsu

## Slide 29: High voltage (~1500 – 2500 V)
Positive (cathode at ground)
  low noise, capacitative coupling
Negative
  Anode at ground (no HV on signal)
No (high) voltage
  Cockcroft-Walton bases

## Slide 30: Housing
NIM A432 (1999) 265

## Slide 31: Compact divider design

## Slide 32: Single photoelectron signal

## Slide 33: PMT single p.e. spectra (noise)
ADC (counts)
ADC (counts)

## Slide 34: Signal for passing tracks

## Slide 35: Energy deposited in scintillator

## Slide 36: Energy deposited in scintillator
Minimum-ionizing peak
pions, electrons

## Slide 37: Energy deposited in scintillator
Minimum-ionizing peak
pions, electrons
Photons and “corner clippers”

## Slide 38: Effect of magnetic field on pmt
Photomultipliers and Accessories
Electron Tubes Ltd (1996)

## Slide 39: Hall D – GlueX detector
Tagger Spectrometer
(Upstream)
Hermetic detection
of charged and
neutral particles in
solenoid magnet
Time-
of-flight
(tof)
Pb-glass detector (Fcal)
Tracking
Cathode strips
Drift chambers
Straw tubes
Superconducting 2 T solenoid
Flux (peak) ~ 2x107 g/s
18,000 FADCs
4,000 pipeline TDCs
50 KHz L1 trigger
600 MB/s to tape
Target
(LH2)
Barrel
Calorimeter
(Bcal)

## Slide 40: Hall D – GlueX detector
Tagger Spectrometer
(Upstream)
Hermetic detection
of charged and
neutral particles in
solenoid magnet
Time-
of-flight
(tof)
Pb-glass detector (Fcal)
Tracking
Cathode strips
Drift chambers
Straw tubes
Superconducting 2 T solenoid
Flux (peak) ~ 2x107 g/s
18,000 FADCs
4,000 pipeline TDCs
50 KHz L1 trigger
600 MB/s to tape
Target
(LH2)
Barrel
Calorimeter
(Bcal)
What to do in a Magnetic Field?

## Slide 41: Some history in photos

## Slide 42: Example, Hamamatsu SiPM

## Slide 43: Hamamatsu S12045 arrays
4x4 array of 3x3 mm2 sensor
Area = 12.7 x 12.7 mm2
57600 pixels
NIM A896 (2018) 24

## Slide 44: Multipixel Limited Geiger mode APD
20-100 µm Geiger APD pixel
Incident photon initiates avalanche
Quenched by resister  device reset
Sum the pixels –> photon counter
Quenching
Resister
Geiger-APD array
Photosensitive surface
RQ
Notes: Each pixel is G-APD – operated in reverse biased mode above some threshold voltage (“breakdown voltage”)
device is metastable – incoming photon can actuate geiger mode avalanche – high gain – quenching resistor uses up charge and brings system back to down to metastable state again
put these G-APD pixels into array connected in parallel (each with own quenching resistor) – as long as photon intensity low enough – essentially count up number of fired pixels to count photons
other slides of interest  31

## Slide 45: Multipixel Limited Geiger mode APD
20-100 µm Geiger APD pixel
Incident photon initiates avalanche
Quenched by resister  device reset
Sum the pixels –> photon counter
Quenching
Resister
Geiger-APD array
Photosensitive surface
RQ
Insensitive to Magnetic Fields!
Notes: Each pixel is G-APD – operated in reverse biased mode above some threshold voltage (“breakdown voltage”)
device is metastable – incoming photon can actuate geiger mode avalanche – high gain – quenching resistor uses up charge and brings system back to down to metastable state again
put these G-APD pixels into array connected in parallel (each with own quenching resistor) – as long as photon intensity low enough – essentially count up number of fired pixels to count photons
other slides of interest  31

## Slide 46: SiPM (3x3 mm2) single p.e. spectra
Dark Counts
Filter=1%
Filter=2%
Filter=4%
Filter=6%
Pedestal

## Slide 47: Measure the Flight Time between two Scintillators
400 cm
100 cm
300 cm
20 cm
Disc
Disc
TDC
Start
Stop
Particle Trajectory
450 ns

## Slide 48: Formalism: Measure time and position
Mean is independent of position!

## Slide 49: Experiment basics
B field ~ 5/3 T
R = 3m
L = ½ p R = 4.71 m
p = 0.3 B R = 1.5 GeV/c
tp = L/bpc = 15.77 ns
tK = L/bKc = 16.53 ns
DtpK = 0.76 ns
bp = p/√p2+mp2 = 0.9957
bK = p/√p2+mK2 = 0.9496

## Slide 50: Particle ID by “tof”: b vs p plot
p+
K+
p

## Slide 51: Summary
  Scintillation counters have a few simple components
  Trigger and time-of-flight systems are built out of these counters
  Fast response allows for accurate timing
  New solid-state light sensors are now available, which allow use in high magnetic fields
  Combining the timing information from scintillator detectors with momentum measurements from tracking detectors, one can determine the mass of passing particles.

## Slide 52: Backup slides

## Slide 53: Electronics
trigger
dynode
Measure time
Measure pulse height
anode

## Slide 54: Formalism: Measure energy loss
Geometric mean is independent of position!

## Slide 55: Example: Kaon mass resolution by TOF
For a flight path of d = 500 cm
Assume experimental resolutions of