## Slide 1
Test results for the recent SIC PWO crystals
Carlos Muñoz Camacho, IPN-Orsay
for the NPS Collaboration at Jefferson Lab
SICCAS, Shanghai, July 23 (2018)

## Slide 2
PbWO4 crystal specification |  |  |  | 
Parameter | Unit | NPS Required | COMCAL/FCAL | PANDA specifications
Light Yield at RT | phe/MeV | ≥15 | ≥9.5 | ≥16
LY uniformity between the blocks (%) |  | <10 (<20) |  | 
LY(100ns)/LY(1us) | % | >90 | >90 | >90
Longitudinal Transmission at |  |  |  | 
λ = 360nm | % | ≥35 (≥25) | ≥10 | ≥35
λ = 420nm | % | ≥60 | ≥55 | ≥60
λ = 620nm | % | ≥70 | ≥65 | ≥70
Transverse Transmission and LY uniformity along the crystal | % | 10 |  | 
Inhomogeneity of Transverse Transmission |  |  |  | 
D λ at T=50% | nm≤ | ≤5 | ≤6 | ≤3
Induced irradiation absorption coefficient dk at λ =420nm and RT, for integral dose >10Gy | m-1 | <1.1 (<1.5) | <1.5 | ≤1.1
Mean value of dk | m-1 | ≤0.75 |  | ≤0.75
Tolerance in Length | µm | ≤±100 | +0/, -100 | ≤±50
Tolerance in sides | µm | ≤±50 | +300, -0. | ≤±50
Surface polished, roughness Ra | µm | ≤0.02 |  | 
Tolerance in Rectangularity | degree | ≤0.1 |  | ≤0.01
Mo contamination | ppm | <10 |  | <1
La, Y, Nb, Lu contamination | ppm | <40ppm (≤100) |  | ≤40
Red denotes relaxed parameters in the 2018 procurement
Prefer a method of labeling crystals that does not leave traces of glue when removing sticker
NPS: 2.05cm x 2.05cm x 20cm

## Slide 3
Dimension measurements
Mitutoyo height gage tool (accuracy ~1um)
Granite table AA class (flatness accuracy ~ 6 um)
Safe and high precision method(~25 um)
Six measurements per one crystal
out-layers

## Slide 4
Optical transmittance measurements: PerkinElmer Lambda 850
Double beam spectrometer
Integrating sphere
Excellent (<0.1%) reproducibility of measurements

## Slide 5
Light yield measurements: Cs-source in a temperature controlled chamber
Temperature: 18 deg C
Light yield measurements: Cs-source in a temperature controlled chamber

## Slide 6
Radiation hardness measurements: high activity 60Co source
30 Gy total dose, at a rate of ~1Gy/min

## Slide 7
Experimental Investigations
Initial Inspection and Labeling
Visual inspection including green laser
Longitudinal and transverse dimensions
Optical transmission
Longitudinal and transverse
Luminescence yield, temperature dependence, and decay kinetics
Radiation resistance
Chemical and surface analysis, as needed

## Slide 8
Conditions for crystal certification and acceptance
NPS collaboration is responsible for the certification of the crystals delivered by SICCAS
Result of certification summary includes these parameters
Individual crystal number
Dimension of crystal
Visual inspection report
Optical properties of crystal
Scintillation yield
Radiation hardness (if measured)
Examples of additional items in 2018 procurement to ensure quality:
  Scintillation kinetics (if measured)
  Non-uniformity of transmission and light yield (if measured)
  Contamination (if measured)
NPS collaboration is not obliged to disclose the results of the certification to SICCAS, except for rejected crystals

## Slide 9
Visual Inspection
Excessive number of small chips along edges or large chips on corners
Long scratches and cracks

## Slide 10
Visual Inspection – cont.
Excessive number of bubbles in bulk
Chemical film on surface
Traces of old labels or markings on surface

## Slide 11
BTCP
SICCAS
CRYTUR
Typical crystal surface quality
Defects result in high, but non-uniform light yield
Looking deeper into defects:
Bubble
Deep scratches
Pits
Surface analysis

## Slide 12
Longitudinal Transmittance:
Correlated with observed defects

## Slide 13
Longitudinal Transmittance:
Concern is the width of the distribution at each wavelength

## Slide 14
Light Yield:
Large variations in light yield
High light yield, but non-uniform due to defects, e.g. bubbles
Very high light yields due to defects in crystal bulk
Non-uniformity of light yield within one crystals due to defects
Concerns:

## Slide 15
Radiation Resistance

## Slide 16
Summary 2017: Three major failure categories
10% fail: LY @420nm < 60% or light yield
13% fail: Bubbles and other defects
8% fail: Chips, scratches, old labels
Out of a subset of ~200 crystals tested, the “bad” crystals fall into three major categories:

## Slide 17
Example report from SICCAS 2017
Measurement methods used?
CX | NPS T360 | NPS T420 | NPS LY100
S021 | 26.7 | 61.2 | 14.3
S022 | 25.6 | 58.7 | 13.3
S023 | 27.2 | 62.9 | 16.0
S024 | 39.8 | 63.1 | 19.7
S025 | 33.8 | 62.5 | 16.0
S026 | 29.5 | 62.1 | 15.8
S027 | 27.3 | 63.8 | 17.0
S028 | 33.7 | 62.8 | 17.8
S029 | 40.9 | 63.5 | 22.7
S030 | 21.7 | 62.4 | 14.2
S031 | 25.06 | 63.47 | 12.45
S032 | 20.3 | 62.8 | 11.5
S033 | 23.6 | 63.6 | 12.3
S034 | 29.2 | 60.7 | 16.1
S035 | 26.8 | 66.2 | 13.8
S036 | 28.8 | 60.6 | 11.1
SICCAS results
NPS results

## Slide 18
SIC and Jlab measurements comparaison (~200 samples)
No clear correlation of results
Transmittance and LY generally lower at JLab than SIC
Should discuss methods and measurement procedures