## Slide 1: SHMS Status HallA/C users meeting
Paul Brindza
Jan 20,2017

## Slide 2: Hall C SHMS

## Slide 3: Overall SHMS Status
All SC magnets are at JLAB and installed
  HB, Q1 tested and accepted
  Q2 accepted this week
  Q3,D TBD
  Experiment Readiness Review was a success
All detectors have been installed.
  Noble gas Cerenkov removed to permit other work
Beamline installation
  DS more than 50%
  US 98%
  Dump upgrade 98%

## Slide 4: Noble Gas Cerenkov Installed Nov 2016

## Slide 5: SC magnets
HB Accepted 2016
Q1 Accepted 2015
Q2 Acceptance status
  Acceptance tests complete
  Reached 3866 Amps > 3660 Amps 11GeV/c after training
  Ramp to 3660 Amps in 30 minutes, hold 1 hour and ramp down
  Cool down in 8 days < 10 days
  Heat leak is ~ 60 watts > 40 watts

## Slide 6: SHMS Q1 and HB Status
Q1:
  Installed on SHMS
  Cooled down
  Power test to 3000 Amps or 110% of operational current
  Accepted 2015
HB:
  Installed on SHMS
  Cooled down
  Power test to 4000 Amps or 102% of operational current
  Accepted 2016

## Slide 7: Q2 entering Hall C Oct 6, 2016

## Slide 8: Q2 Has arrived Oct 6, 2016

## Slide 9: Q2 installation Oct 7, 2017

## Slide 10: Q2 cool down Nov 11, 2016

## Slide 11: Q2 first helium fill Nov 19 , 2106

## Slide 12: Q2 training curve
Chart: Q2 Quench Current
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
current | 3316 | 3446 | 3542 | 3602 | 3603 | 3716 | 3744 | 3773 | 3804 | 3824 | 3837 | 3822 | 3835 | 3851 | 3844 | 3866 | 3847

## Slide 13: Q2 beam side Jan 19, 2017

## Slide 14: SC magnet part 2
Q3
  installed and assembled
  Connections underway to cryo, power and controls
  Passed leak and pressure and hipot tests
  JT valves, pressure sensors and thermometers calibrated
  Set up of voltage taps and quench detector TBD
  I&C checkout TBD
  Purification and cooldown TBD
  Power tests TBD

## Slide 15: Q3 in Hall C Dec 7, 2016

## Slide 16: Q3 Installed Dec 9, 2016

## Slide 17: Q3 CCR Jan 18, 2017

## Slide 18: Q3 CCR Jan 19, 2017

## Slide 19: Q3 from beam side Jan 19,2017

## Slide 20: SC magnets part III
Dipole
  installed
  Chimney assembly underway
  Connections to cryo, power and controls underway
  Hipot test passed
  Pressure sensors and thermometers calibration done
  Leak and pressure tests done
  JT valves calibration underway
  Set up of voltage taps and quench detector TBD
  I&C checkout TBD
  Purification and cooldown TBD
  Power tests TBD

## Slide 21: Dipole installation on SHMS Dec 2016

## Slide 22: Dipole CCR work Jan 18,2917

## Slide 23: Dipole CCR Jan 19, 2017

## Slide 24: Dipole beam side Jan 19, 2017

## Slide 25: What’s left besides magnets?
SHMS rotation
  SHMS can be rotated under power by an operator
  Mount and wire angle limit sensors
  Test rotation in full automatic mode
Assemble SHMS vacuum system
  Align magnets to a common center
  Magnet to magnet bellows
  Shutter and vacuum window
  Install pumps and test under vacuum

## Slide 26: What’s left besides magnets II
Down stream beamline
  Complete tube installation ~ 50% now
  Install pump and test under vacuum
Up stream beam line
  One short piece to install between scattering chamber and existing beam line

## Slide 27: Hall-C Project Status
HB Magnet, Sieve Slits, Scattering Chamber
Beamline Vacuum Pipe Sections
Drift Chambers
Installed In SHMS
Compton
Polarimeter
Noble Gas Cerenkov
Moller Quad Power Connections

## Slide 28: Conclusions
Q3 and Dipole are in a race to cool down first
Both should be cold by end of February
Magnet operations for the key Performance Parameter validation should not be difficult due to 6.3 Gev/c electron beam
Magnet 6 GeV/c (11 GeV/c))
  Dipole 1854 Amps (3400 A)
  Q2 1996 Amps (3660 A)
  Q3 1362 Amps (2500 A)
  Q1 1636 Amps (3000 A)
  HB 2181 Amps (4000 A)

## Slide 29: Old friend HMS 25 years old! Now at 7.4 GeV/c

## Slide 31: Hall-C Project Status
Q2 in Hall C and On the SHMS
Dipole Left Sigmaphi
23-SEP
Final Welding of Q3