## Slide 1: Tests of CAEN 1190 Multi-hit TDCs
Simona Malace
Brad Sawatzky and Brian Moffit
JLab
Hall C Summer Workshop Aug. 15-16 2013, JLab

## Slide 2
Scope
 CAEN 1190 multi-hit TDCs will be implemented in Hall C for the 12 GeV running in both HMS and SHMS
 In an initial bench test we studied the high rate behavior/capabilities of TDC1190
Outline
 General remarks on TDC1190
 Experimental setup
 Data Structure/ Data Decoding/ Data Analysis
 Results: resolution, hit maps, measurement histograms
 Still to do…

## Slide 3
General Remarks on TDC1190
CAEN 1190 multi-hit TDC: 128 independent multi-hit/multi-event
time-to-digital conversion channels organized in 4 High Performance TDC (HPTDC) chips (32 channels each)
The data acquisition can be programmed in “EVENTS” (“TRIGGER MATCHING MODE” with a programmable time window) or in “CONTINOUS STORAGE MODE”
It can be enabled to the detection of the rising and/or falling edges of the pulse width
The board houses a 32 kwords deep output buffer that can be readout via VME as single data, block transfer or chained block transfer in an independent way of the acquisition itself

## Slide 4
General Remarks on TDC1190
TDC chip architecture
40 MHz input clock but 100 ps resolution can be achieved with PLL multiplication
Each channel can store up to 4 hit measurements before being written into a 256 words deep L1 buffer shared by 8 channels
In trigger matching mode a function will match hits with triggers based on trigger information stored in the Trigger Fifo (bunch id and event id) and a programmable matching window
 The read-out FIFO is 256 words deep and enables data to be read out while others are being processed

## Slide 5
General Remarks on TDC1190
Power ON: Default
For our test:
Trigger matching: enabled
Matching window width: 900 ns
Matching window offset: -1200 ns
Subtraction trigger time: enabled  time measurements are referred to the beginning of the match window
We read the TDC1190 via Block Transfer, Direct Memory Access (BLT DMA)
Notes: Lock Transfer

## Slide 6
Experimental Setup
-1200 ns
Beginning of match window
match window width = 900 ns
-560 ns
-300 ns
-960 ns
This 1 kHz reference pulse goes in select channels:
0, 15, 16, 31 in TDC1
32, 47, 48, 63 in TDC2
64, 79, 80, 95 in TDC3
96, 111, 112, 127 in TDC4
Common stop
This 1 kHz reference pulse is OR-ed with a variable rate pulse (mimics noise) and goes in all channels
Ref. pulses: delayed copies common stop

## Slide 7
Data Structure/Data Decoding/Data Analysis
dummy
filler
Global header
TDC1
header
TDC1
data
TDC1
trailer
Global trailer

## Slide 8
Data Structure/Data Decoding/Data Analysis
We use the Hall A analyzer; the THaCodaDecoder has been modified to extract information from: global/chip headers and trailers and form measurement words
Number of hits
Measurement for each hit

## Slide 9
Results: Hit Maps
Rate per channel: 1 kHz, 2kHz in select channels (noise = 1mHz)
At very low noise rate by construction we expect 2 hits in select channels 0, 15, 16, 31, 32, 47, 48, 63, 64, 79, 80, 95, 96, 111, 112, 127 and 1 hit only in all the rest

## Slide 10
Results: Hit Maps
Rate per channel: 501 kHz, 502 kHz in select channels (noise = 500 kHz)
At 500 kHz noise rate there is a large probability (~45%) to have 3 hits in select channels 0, 15, 16, 31, 32, 47, 48, 63, 64, 79, 80, 95, 96, 111, 112, 127

## Slide 11
Results: Hit Maps
Rate per channel: 1.1 MHz, 1.2 MHz in select channels (noise = 1 MHz)
At 1 MHz noise rate there is an even larger probability (~90%) to have 3 hits in select channels 0, 15, 16, 31, 32, 47, 48, 63, 64, 79, 80, 95, 96, 111, 112, 127

## Slide 12
Results: Hit Maps
Rate per channel: 5.1 MHz, 5.2 MHz in select channels (noise = 5 MHz)
At 5 MHz noise rate it fails; there should never be events with 0 hits

## Slide 13
Results: Raw Measurements
Rate per channel: 1 kHz, 2kHz in select channels (noise = 1mHz)
2 reference pulses at 1 kHz rate
I picked this one for subtraction
Measurements are reported w.r.t. the start of the match window
this
farther away from common stop than
this

## Slide 14
Reference Pulse Subtracted Measurements
Rate per channel: 1 kHz, 2kHz in select channels (noise = 1mHz)
Resolution
0.127 – 0.14 ns

## Slide 15
Reference Pulse Subtracted Measurements
Rate per channel: 1 kHz, 2kHz in select channels (noise = 1mHz)
Here the chosen reference pulse has been subtracted from all other measurements (ref_pulse - measurement)
Chosen reference pulse
The other reference pulse

## Slide 16
Reference Pulse Subtracted Measurements
Rate per channel: 501 kHz, 502 kHz in select channels (noise = 500 kHz)
Here the chosen reference pulse has been subtracted from all other measurements (ref_pulse - measurement)
Chosen reference pulse
The other reference pulse
The gap is probably generated by the noise and reference pulse “walking” on each other:
the leading edge of the noise pulse comes first and there is overlap with the ref. pulse a gap on the + side
the leading edge of the ref. pulse comes first and there is overlap with the noise pulse a gap on the - side

## Slide 17
Reference Pulse Subtracted Measurements
Rate per channel: 1.1 MHz, 1.2 MHz in select channels (noise = 1 MHz)
Here the chosen reference pulse has been subtracted from all other measurements (ref_pulse - measurement)
Chosen reference pulse
The other reference pulse
At ~ 1MHz per channel, ~128 MHz integrated rate still looks O.K.

## Slide 18
Reference Pulse Subtracted Measurements
Rate per channel: 2.1 MHz, 2.2 MHz in select channels (noise = 2 MHz)
Here the chosen reference pulse has been subtracted from all other measurements (ref_pulse - measurement)
Chosen reference pulse
The other reference pulse
At ~ 2MHz per channel, ~256 MHz integrated rate it appears to fail

## Slide 19
Summary
 We tested the high-rate behavior of CAEN 1190 multi-hit TDCs:
 it appears to function properly up to ~ 1 MHz per channel, 128 MHz integrated rate
 we programmed the TDC for 100 ps resolution; the measured resolution is 120 -140 ps
 Still to do:
 sanity checks …. Again!
 simulate a “hot wire” input and verify if it fails in a similar fashion as shown here

## Slide 20
Unrelated: the maPMT H8500C-03 test results I presented at the last Hall C meeting have been organized into a paper that was accepted this week for publication in the Journal of Instrumentation