# MEMORANDUM

To:		Distribution

From:		F. Dylla

Subject:	FEL Upgrade Project Weekly Brief  - June 6-10, 2005

Date:		June 10, 2005

Highlights:

We spent two day this week on machine set-up for high power 2.8 micron operation and devoted the latter half of the week to low-power, broad-band operation (1.1-1.8 micron) operation for the second user run for Harvard Mass General Hospital (MGH). This run was very successful according to our Harvard collaborators.  Data was gathered on the differential heating of fat vs. non fat tissues  at wavelengths that correspond to C-H stretch overtones at 1.2 and 1.7 microns.  The broad-band mirror assembly allowed the entire resonant regions to be scanned at 10 nm increments.  Kudos were given to our Optics and I&C teams who arranged for local control of the FEL and local viewing IR videography of the exposed samples in User Lab 1. These feature greatly facilitated the experiments.

After the MGH experiments were completed late on Friday afternoon, we set up a follow-up series of spot size calibrations in the Dahlgren set-up in User Lab 1 as requested by our Dahlgren colleagues who collected valuable data early last month on IR sensor damage.

Next week, FEL operations will be devoted to high power operations at 2.8 microns.

By all accounts the workshop we co-hosted this week (June 7-8) on the “Free Electron Laser Applications to Biology and Medicine” was successful.  The workshop attendance was limited to 60 attendees from the 7 sponsoring institutions and others who indicated interest.   Four program areas were identified as having high potential for FEL applications: studies of phototherapies, UV photo-damage effects, high brightness imaging from THz to UV, and laser based diagnostics.   A workshop summary and program plan is being written by topical chairs nominated by the attendees.

Progress on the UV FEL system this week included a successful mid-course design review on the UV cavity assemblies and the completion of the first of the THz chicane dipoles (GW’s) at the vendor.

Management:

We chaired and shared in the organization of this week’s laser bioscience workshop.  As noted above, the workshop was successful and generated quite a bit of favorable notice for the capabilities of the FEL Upgrade.

Effort was devoted to contributing FEL project data into the Lab’s annual update of our overall staffing plan as we develop detailed budget scenarios for FY06.  We also contributed strategic planning documentation for the lab’s 5 year business plan as requested by DoE’s Office of Science.

Operations:

Monday was spent in recovering from the repairs to the buncher RF system.  The repairs were quite effective in stabilizing the buncher operation.  The injector was phased and we steered up the beam in the wiggler.  We spent some time trying to suppress the dispersion in the first arc.   After managing to suppress the dispersion pretty well we realized that the procedure was flawed.  We will have to try again next week. 
   On Tuesday we set up the machine for lasing at 115 MeV in the 1.1 to 1.8 micron wavelength band.  The laser lased quite strongly even at 1.2 microns.  We spent some time on Wednesday calibrating the setup for MGH and testing out single shot mode.  We also tried to use the new

88 MeV injector setup but were unable to get a good match to the wiggler.  Since the performance was quite good with the original injector setup we stuck with that one.  We had one problem with an emitter turning on in zone 3.  The gradients were reallocated to reduce the RF heat load in that zone. 
   Thursday and Friday were dedicated to beam delivery for the MGH  experiment studying the effects of light in the 1.1 to 1.8 micron  range on lipid rich tissues.  The machine range quite stably and well.  We could get over 4 J in a 100 msec pulse even at 1.2 microns.  Running with short pulses and 1.25 mA we could get up to 170 W during a macropulse delivered to user lab 1 using the broad-band mirror set.

WBS 4 (Injector):

More high voltage tests were performed this week on the flat electrode coated with a thicker field emission suppression film. At 25 MV/m the average field emission current is around 90 pA. The highest gradient in the gun support tube electrode is 12.4 MV/m at 500 kV operation. More conditioning and testing is planned for next week. 

A POISSON model of the new gun chamber shows that the gradient relaxes slightly on the support tube near the ground ring. This relaxation is beneficial and does not affect the gradient at the cathode. 

This week we started to work on drawings for the coil that will be used for NEG coating the new gun chamber for the Injector Test Stand. The tube segment polished earlier has been ultrasonic cleaned and sent to the machine shop for welding the ring that supports the ball cathode. Polishing of the second tube segment continued this week, the transition from silicon carbide paper to diamond paste is just starting. Equipment is being moved out of FEL lab 2 to our new lab in the ARC building as part of the preparations of the labs for FEL users.

With the help of Daniel Mihalcea and Court Bohn (Northern Illinois University), Kevin Beard has begun to run large simulations in their Linux cluster at FNAL both of the space charge effects in the FEL injector (with PARMELA) and to estimate the beam breakup (BBU) current thresholds for the FEL (with MATBBU) when the FEL03 cryomodule is replaced with the “Renascence” cryomodule. Testing is under way and we expect preliminary results next week.

The new FEL injector configuration with a stronger solenoid setting (the 88 MeV configuration) was modeled this week in PARMELA. Preliminary results show very good agreement between code predictions machine behavior. Beam-based measurements are being planned for next week.

Gun HVPS - Operational in all respects.

WBS 5 (SRF):

The first cavity is going into the VTA this week for initial testing and qualification.  The second cavity is positioned in a test stand and will follow next week.

Preliminary beam line bellows transition designs for the third harmonic cavity have been forwarded from AES to JLAB for incorporation into the injector unit layout.

AES has completed the purchase order for the titanium helium pressure vessels; delivery is expected this fiscal year.

The requirement for HOM damping and power coupling has been identified as an area needing further clarification.  This will be discussed during the coordination meeting.

A date for the technical coordination meeting between AES and JLab has been set for 21 June 2005 at the JLAB site.

WBS 6 (RF):

RF - All systems are operational.  This week the Buncher's RF Control Module was found to have corrupted calibration files and its pre-klystron RF amplifier was found to have several loose internal cable connections.  After making these corrections, the Buncher phase drift errors are at 1 degree p-p over several days and the large 5 to 10 degree phase jumps are gone.

WBS 8 (Instrumentation):

In general, the week has been dominated by support of Ops for LOP-05-004 (MGH, Rox Anderson).  The I&C set up time for the experiment was much quicker this go around due to the new A/V cart.  This provided us with a much more centralized setup which cleaned up cabling and reduced trip hazards.  The cart comes equipped with a video monitor, two quads, a vcr, a computer interface, a microphone input, and a video chassis. All of the cameras on the test equipment were connected to the system along with the area monitor in lab 1. The experiment seems to be moving along just fine. 
   Our efforts to build a rack mount PC to host the Acqiris AP201 card have come to a halt due to a PCI incompatibility we have discovered between the A2D card and the PC-motherboard. We are looking into possible BIOS issues versus using another motherboard.   A 150 MHz Bandpass filter evaluation board was designed this week.  The parts have been ordered to fabricate the board and the quote for manufacturing of the PCB has been received.  EECAD has completed all the documentation for the project and it has all been signed off on.  Progress continues on the 6kW polarity reversing switch.  Currently several layout ideas are being examined to determine the best package design for the system.  Parts continue to arrive for the beam viewer crate upgrade.  The back plane design for the crate is nearing completion and will be turned into EECAD at that time.  Pre commissioning of another stepper motor chassis was done this week. Testing will begin once the move in to the new trailer is complete.  Wiring and debugging of the 4 Channel Solenoid Fail-safe Interlock prototype PCB has been completed.  During debugging, it was found that the "In_Limit_Status" protection section of each channel was erroneously checking the status of its corresponding 'in' limit in addition to the opposing three solenoids'.  Corrections were made so that each channel only checks the status of the remaining three channels' in limits.  All changes were documented and submitted to EECAD.   Discussion has begun on the ceramic viewer in 3F12. After installing Auto CAD 2005 on the computer in Area 51, we found the old mechanical drawings made by Jessica Sheehan for the viewer. We know that the problem is in alignment of the mirrors, but these drawings will be used as a basis for the corrections that need to be made so that the exact problem can be identified. Real work on it will not start until the vault opens June 20th, where the plan of action will probably be to disassemble it and begin a major overhaul in the alignment mechanics. The final items in the DMA processing lab were completed early this week. The area camera needed to be connected to its respective monitor in the control room. We will align this camera again today and possibly, if need be, put some locktite on some of the bolts/screws. 
   All of the parts for additional planned temperature monitoring crates (using the Silicon Diodes) are on order and assembly of the boards should begin early next week.  Three crates have been assembled and are ready for the boards and testing.  The RTD portion of the design is being reviewed and adjustments made to achieve the desired results.   The new design of the magnet trim cards is progressing nicely.  The prototype is complete and testing and fine tuning is in progress.  Programming issues have been resolved and the next step is load tests and temperature testing. Bids have been received from three vendors for the production and population of all of the components involved.  Our single-shot trigger archiver has worked well this week. The software tool automates the task of chronicling the laser pulse parameters for the users and allowing a time stamped and record of the entire experiment to be printed out and handed to the user as a part of their work here. This system will continue to evolve into a system performance tool. Currently we have time stamp faults from the RF system and gather the relevant EPICS information with the fault. 
  The process of moving into the trailers has taken full swing and clearing out Lab 5 and Area 51 are well under way.  The two transportainers have been reorganized and the majority of the parts in area 51 have been split between them and the new trailer.  Other non-essential items such as spare pcb's have been removed from lab 5 and placed in the trailer.  Additional parts for shelves have been ordered, and the move in will continue today and into next week.

WBS 9 (Beam Transport):

Sextupole (SF)

• Fabrication continues and assembly and potting started at New England Techni-Coil. All 66 coils are wound. The concept of ramp blocks at transitions between layers of these two-in-hand wound coils was successful at eliminating shorts. Machining/grinding of all core pieces is complete. The first core was assembled and inspected, meeting specification. The multiple, accurate “T” & Hexagon concept works to create an accurate core from 12 parts that can be assembled around the coil. Coil molds are complete and a first cycle of potting 3 to 6 per day will start tomorrow. Ceramic tubes that act to isolate the LCW circuit from the coil are starting to loom as a missing item. The vendor we recommended has not come through. Jlab’s Larry Munk is working the issue. The stands Jlab supplied to mount the magnets on stands have a small (.005 inch) interference problem that NETC will machine to a custom fit. Also, some tapped holes for brackets in the core were miss-specified and need to be re-done. NETC plans shipping all units the week of July 4.

Replacement Chicane Dipoles (GW)

•  Fabrication continues and assembly of one unit is complete at New England Techni-Coil. The first dipole assembly is good. The Purcell gap glue-down of the second core meets specification – this potentially risky task, for a vendor doing it for the first-time, is not risky for NETC. The sixth coil (of 8) is almost wound. A cycle of potting two coils per week should start next week. Machining of the top and bottom pole plates of the third dipole is 75% complete. Shipment of all units is anticipated the Week of July 4.

UV Line

• The most important quality of the QX Quadrupoles of the UV Line is that they meet 1 part in 1000 set ability and 1 part in 10,000 reproducibility specifications. Magnet Test made substantial progress toward that goal this week. They spent the week investigating only reproducibility, powering a typical QX using both the standard Bang-Bang Hysteresis and Setting with Undershoot as well as the proposed Slow, 3-part Linear Ramp with Setting from the Particular Polarity Protocol. They used all the available Power supplies: BOSS, Danfysik, Trim Card with EPICS control as well as the new, Trim Card II, (“power end” only). With all readings of field gradient integral from a rotating coil, they also took readings of an accurate hall probe glued to the pole tip. They found that the field, as measured by the hall probe, when corrected to the current read by the transducer resolved at 1 part in 10,000 while the rotating coil resolved field gradient integral at only 5 parts in 1000. Thus, they are taking the hall probe as the better instrument for investigations. They find both the Bang-Bang and the Slow Ramp Protocols are reproducible within 1 part in 10,000 using the BOSS and Trim Cards. 

They are now starting investigating set ability using the same cast of resources. We learned from our FEL test that actual usage of quads always involved iterating to a setting (while looking at a beam spot) between forced, double hysteresis loops. This method produced different fields (on a hall probe on a pole tip) in out test from those resulting from just decreasing the current from the top of a positive hysteresis loop. The set ability tests will compare new field maps generated by the hall probe glued to a pole tip using the original top down method and the map generated at settings after local hysteresis loops. They plan to test this QX Quadrupole in the Magnet Test using the Epics Controlled Trim Card Power Supply and then transfer the quad to the FEL Gallery. There, the same trim card will run the magnet in one of the Trim Racks and using the hall probe, verify the “between buildings” set ability and reproducibility. This Gallery Test also allows us to address set ability and reproducibility under scenarios where other trim cards are working and cycling. 

This next item is “Huge”. This week, Ken Baggett (on loan from the Alignment Group) has finally been able to make an improvement to the Test Stand that was on a wish list since the dawn of Magnet Test. Using a piece of software from Christiana Grenoble, he obtained semiautomatic, linked EPICS control of the available trim card power supply (in the Injector Test Cave) and the functions of the rotating coil test stand. This means we can test magnets with the same power supply system that powers and controls them in CEBAF and FEL and put aside the cranky BOSS supply. He predicts this semiautomatic ability will be fully functional next week.

• The Region Assembly Drawing of the 6Fsub1 region of the UV Line was signed off. Girder drawings for that region are still in preparation. Regional Assembly of the 8Fsub4 (Last) is also in preparation but because of its similarity to 6Fsub1, the stand alignment positions are released to the Alignment Group so they can use them during the June Down Period.

• The 8Fsub2 (GW Chicane) Sub Region Assembly Drawing and chamber drawn by Advanced Energy Systems remains in signature while the stands are signed off. Note that the Jlab Shop has started the chamber procurement.

UV Wiggler Progress

•	No progress on the new Wiggler vacuum chamber.

IR PM Wiggler Vacuum Chamber Progress

•	The raw tubing for the chamber was ordered.

WBS 10 (Wiggler):

STI Optronics is making good progress on the permanent magnet wiggler.  The figures below show the finished frame before painting, the control panel, and the motors and micrometers for the gap mechanism.  We ordered the tubing for the vacuum chamber for this device this week. 
We have decided to hold the kickoff design review for the UV wiggler next Friday via video-conference.  We can afford to do this since most of the design details have already been worked out for the infrared wiggler.

WBS 11 (Optics):

FEL mirrors:
   This week we completed the debonding of the 1.6 micron OC, recheck of it's figure, and clean up of the holder.  It goes back into the braze furnace on Monday.  We were unable to get one of our 2" holders to pass leak check, after multiple braze attempts.  Our spare is ready to receive the optic.  Our recently bonded HRs have low astigmatism when undeformed, but are showing rapidly increasing astigmatism when deformed, so we are looking at debonding them to determine what is happening.  I suspect we have an adherence issue, which we can quickly resolve.  We were informed by the vendor that one 1.6 micron HR will ship today, while the other, and our 1.06 micron, 96% R OCs will ship next week.  This is still sufficiently early to allow us to meet the installation schedule for the last part of the month.

As mentioned in the Operations section, our broadband mirror set (hole outcoupled) worked sufficiently well to obtain very good data.

Other activities:
   We are juggling a number of activities this week.  We supported operations for the Mass. Gen team.  It appears that our beam quality measurements will be observed by the Dahlgren team.  They may wish to do a short set of calibrations as well.  We will be measuring the near-field spot sizes two ways, with a knife edge test, and with a scanning lens and our PyroCAM.  We did a brief set of measurements to check out the latter technique earlier this week.  We also began the move of some of the contents of UL 2, our "home" for some 5 years, over to the new trailer.  This is freeing up room so we can start building the hutch for the Dahlgren team.  We welcome Luke Osborne, a student from the College of William and Mary, to the team for the summer.

UV FEL
   We received the LabView drivers for a fiber-based spectrograph.  Initial checkout shows they give us the entire spectrum (4096 pts), while we usually need only a small region.  So, we'll look at the program to see if we can modify it appropriately, to all it to be usefully accessed with our EPICS controls.  We had a mini-review of the optical cavity.  It generated a list of comments, and two issues.  The overall design is judged to be sound, and we are having the mirror mounting internals for one assembly fabricated.  It should be complete by July 1.

We removed the internal components of the mirror cassettes in User Labs 2 & 4.  They will be modified for better vacuum practice in general, and for high average power (10 kW) operation in the case of the UL2 mirror cassette.

Terahertz Project:

The hutch in lab 3 is ready for certification.  Specifications for the vacuum chamber for the pump-probe measurements continue.  The THz spectrometer is working again.