# MEMORANDUM

To:		Distribution

From:		F. Dylla/grn

Subject:	FEL Upgrade Project Weekly Brief  -  November 13-17, 2006

Date:		November 17, 2006

Highlights:

During the first part of this week we removed the cryocooling from the 1.6 micron mirror and re-installed our standard water-cooled mirrors for 1, 1.2, 1.6 and 2.8 microns in the outcoupler can. (Note that our current version of the cryocooling option does not allow multiple mirrors to be accessed in the optical cavities).  This return to our standard configuration will allow us to continue our FEL studies for JTO and Navy programs over the next month and a half.  As discussed below in the Operations section after re-establishing FEL operations, we initiated detailed studies on short Rayleigh range operation.  These tests will continue through next week.

We also held a successful design review for our 750 MHz 1 Amp cryomodules project.  Although there are still some outstanding design issues to resolve, overall the effort is in good shape to proceed with element design engineering and component development tests. This design review was a key deliverable in our Navy program plan for FY07

Management:

We prepared and submitted follow-on Phase II proposals for our current JTO projects and completed our submission to SAIC for our contribution to their follow-on.  Those proposals will be defended at the JTO office in 1 1/2 weeks.

A coordination meeting for the Yale/Hampton University LIPPS experiment (the axion search) was held this week to coordinate efforts for their requested engineering and first data runs. The monthly teleconference for our Biomedical Photonics Consortium was held today. Plans were initiated for a 1-day workshop early in the new year to coordinate CY07 proposal development.

Preparations continued for on-site meetings with NAVSEA and ONR on December 1 and 8.  ONR distributed the final agenda for our Jan.17-18 annual review to the review panel.

Gwyn Williams presented an invited talk on how the JLab FEL maps into the landscape of 4th generation light sources, and Fred Dylla presented an invited talk on the development of high voltage from Ben Franklin to the International Linear Collider at this week’s AVS Symposium in San Francisco. Marcy Stutzman from the CEBAF Source group presented a comprehensive review for the need for extreme high vacuum in photoemission electron sources. Many of our FEL users attended this meeting.

Operations:

This was a fairly short week since we spent much of the week pulling out the cryomirror and putting the room temperature mirrors in place.  We were able to lock up late Wednesday and check the injector drive laser alignment and turn on RF.   We turned on Thursday morning at 2.8 microns, switched to 0.94 micron operation for JTO tests, and tuned up the laser at that wavelength.  With the tuning almost done the CHL tripped.  It took most of the day to pump down the CHL.   During this time we calibrated some diagnostics for the JTO work.  We continue the JTO studies of short Rayleigh range today.  The efforts today are to carefully calibrate the Rayleigh range measurement using several different techniques.

WBS 4 (Injector):

The photocathode lifetime is around 1667 Coulombs after adding up the extracted charge from last Friday's run. So far the photocathode has delivered 472 Coulombs since Nov. 1 without re-cesiation. With about 2.5% QE left, the Drive Laser polarizer is set at 40% for full charge. 

The operating vacuum conditions in the gun improved significantly after the Helium was pumped out from the injector cryomodule during the shutdown earlier this week.

We have been studying the effect of soaking temperature on the GaAs wafer surface morphology during the heat clean cycles for photocathode activation. There is clear evidence from AFM data between a fresh GaAs sample and a heat-cycled one that the temperature used so far is too high and is causing As depletion and consequently increased surface roughness. This increase in surface roughness leads to drive laser scatter light that reflects off the mirror-finish anode plate and lands on other areas of the photocathode outside the drive laser illuminated spot. We suspect this scattered light is the cause of beam halo observed as the photocathode undergoes increasing number of heat cycles. In a preliminary test, the scattered light from the heat-treated sample is 20 times larger than that of a fresh sample, as measured by the CCD camera. A simple experiment has been setup to find the threshold temperature for the As depletion. We expect to get some results early next week. Then we will make a photocathode using this threshold temperature to make sure it is high enough to produce good QE. 

D. Bullard helped remove the wiggler cooling plates and made some more progress polishing the molybdenum anode plate for the test stand gun. Drawings for the motorized photocathode retractor system to be incorporated to the FEL gun for reducing the FEL down time during re-cesiation are being reviewed. The corona shield fabrication for this new scheme was sent out for bid.

WBS 8 (Instrumentation):

The first part of this week was spent debugging and repairing a set of embedded BPM electronics that Operations had noted as failing.  The electronics were removed and after several series of tests it was found that due to the on board regulators the board was over heating.  To remove the heat and to dissipate some of the load on the regulators some in series diodes were added to reduce the input voltage and also a fan was installed to increase air flow.  This problem is being taken care of on the next revision of the electronics.  In the process of testing the electronics, we also replaced the ADC and recalibrated the electronics.  The calibration code was uploaded and the electronics are working fine.  Other down time activities included repairing SLM2F08 attenuators and installing interlocks for the new windowed valve VBV1V01.  Various cables were also installed since the risk of crashing the machine was eliminated.  The PC felpx02 was repaired and reinstalled in the vault.

The upgrade for the Wiggler control software was tested.  The serial communication to the Wiggler's Galil motor controller is not working properly.  The IOC appears to be sending the commands properly, but the Galil controller is not transmitting the values back to the IOC.  This is currently being investigated.  The QSPI module on the Coldfire board has been programmed, which will control the General purpose board.  We programmed the FPGA code for the new version of Single board IOC.  The second revision layout of the Single Board IOC Module is being populated.

Much time has also been spent this week getting together everything that is needed for the Beam Viewer upgraded crate prototype.  We've verified the operation of the test box to be correct and are in the process of testing the Control Cards.  Some tweaking of the PLD code on the control cards has been done and some further tweaking remains.  The assembly of one of the backplanes has been added to the queue since the board arrived from the board house early this week.  Currently the cabling to break the backpanel out and into the rear field connections panel is being configured as well as the locations for the power supplies.  The latest revision of the Single Board IOC has arrived and the first one of these will begin testing today.  The plan is to use this new revision with the 3U General Purpose I/O(GPI/O) card as the processor for the prototype Beam Viewer crate.  The GPI/O has also been added to the queue for assembly. We are waiting on parts to finish populating the 3U General Purpose I/O card.

We've also taken the time to coordinate with the EES/DC group for the sextupole power supply installation that will take place during the next down.  We're currently acquiring the requested external ADCs needed for their setup as well as getting the GPI/O tested for installation in their crate.  To give them some time to fully test their system into a load we've began the installation of the reversing switches onto the power supplies.  The water line manifold parts are all in hand and these are partially assembled.  The other parts to provide the switch cooling are being acquired.  A required modification to externally interlock the switches to the power supplies is being planned so that it can be implemented as soon as possible.

Progress was made on the LiPPS experiment with the completion of the power lead installation and connections.  Installation of the Klixon wiring and water cooling hoses is in progress.  Planning is in progress for 'safing' the magnets and power supply which includes barriers, flashing beacons and installing additional kill switches in the User Lab.  Additional Beam viewer Interface boxes, PC boards and parts have been ordered In support of the Optics group User Lab upgrades.
    We are fabricating the 74.85 MHz chassis for the FEL Drive Laser.  Assembled the computer numerically controlled milling machine.  In support of the LPSS upgrade for labs 3A and 6 some cables were installed.  Also some cables were installed to support lab user operations.  Work was done on the driver box for the sputter chamber; the box is complete except for a 12V power supply, which has been ordered from Astrodyne and is due in the week after Thanksgiving.

Rich Evans has spent the week in preparation for an oral defense for his Masters Degree in Applied Physics at CNU.   The defense is today at CNU and (if it goes well) will be the conclusion of his graduate studies at CNU.

Work has been done on the calibration procedure of the log-amp base BPM electronics. The proper calibration is one of the critical aspects of this sort of the BPM electronics. First to speed up the calibration procedure, while keeping it precise, a precise 1-to-4 RF splitter has been assembled, calibrated and tested for long term stability. The goal of the BPM electronics calibration is to collect data, which would be sufficient to reconstruct the level of the input RF signal with the accuracy of about 0.1 dBm or better by measuring the output DC voltage. Therefore during the calibration the RF signal level at the input of the electronics has to be knows with the accuracy better that the 0.1 dBm. For the calibration we use an RF generator with the amplitude accuracy better than that. To calibrate all four channels of the electronics simultaneously it was desirable to have 1-to-4 RF splitter with the difference in the attenuation from, channel to channel, much less than the 0.1 dB. Commercially available splitters do not offer this level of precision. By putting adjustable attenuators on the output of a 1-to-4 splitter we could adjust the difference in attenuation between the different channels to be on the order of 0.01 dB or even better. As the next step the long term stability of the combination of the calibration RF generator and the splitter was measured. Driving the test we would drive the splitter with the RF generator and measure the power level at the output of the splitter with broad band RF power meter. That has been done for a time span of about two weeks. The RF level at the output would not change by more that 0.02 dB, which is a combined variation of the RF generator the splitter and the power meter. Hence the variation of the splitter itself is less that that and is very good level of stability for the log-amp based BPM electronics calibration. Use of the precise splitter reduces the time needed for the calibration of one set of the electronics, which is about 1 ( 1.5 hours, by a factor of 3 ( 4.

We also made a transition to another RF generator to be used for the calibration. For that reason the calibration software was rewritten to use the new generator.

WBS 9 (Beam Transport ):

IR Machine Operations

•	We removed the temporary cooling plates on either side of the Wiggler Beam Chamber to enable it to close its jaws to the full extent.  This will bring back the large temperature excursion and deflection of the tube at high beam currents.

Improvements and Upgrades

•	With regard to SF sextupole, Neil Wilson’s installation crew is not able to perform a disassembly and reassembly to mimic the vault’s cramped conditions due to higher priority work.

• 	Facilities Management continues working on raising the capacity of that Drive Laser Room’s cooling unit, a task that will enhance our ability to keep the drive laser phases stable.

•  Concerning work on the second wiggler chamber: we are still about to solder on cooling tubes. We determined what to order for flashing a copper coating to the inside.

• 	In work for the New Gun and injector Test Stand, McAllister Technical Services hasn’t come back with their final design of the proposed device. With regard to the contract to Riggins Co. to bring the tank for the high Voltage Power Supply back to operable condition with a “hydro” test, the tank was delivered to them.  With regard to vacuum bake-out of the new gun chamber, we haven’t had time to restart this effort.  Larry Munk has sent out the bid drawings for the aluminum shell that covers the apparatus at the high voltage end of the ceramic stack for both the injector test stand gun and the new upgrade to the gun in the vault. Ron Lassiter continues working on the final drawings of the apparatus that will allow remote withdrawal of the cathode stack for cesiation without opening the gun tank.

Experimental program

Experiment to see a pseudoscalar particle that couples to photons (LIPPS)

•  	The group was able to insert the new pole tips in the proper location and mount the light tight regeneration tube in the second set of GW dipoles. The stand for the Optical Dump is mounted and the Refrigerator/Camera stand, the mounts for the water-cooled turning mirrors and the remaining parts for magnet covers are being made by the shop.  The GW magnets were wired up by the electricians and they then disconnected and locked out the magnet until the covers can be put on and the SOP written and signed off.  The group is planning for an early December  run.

WBS 11 (Optics):

As noted in the Operations section the first part of this week was devoted to removing the cryomirror and associated hardware and installing the standard water-cooled mirrors.  These are for operation at 2.8, 1.6, 1.2 and 1.06 microns.  The 1.2 micron mirror will be used in conjunction with the broadband HR mirror.  The new 1.06 micron OC mirror is of a new coating design that should allow harmonic lasing at 355 nm as well.  This will dovetail neatly with determinations of the accelerator's performance.  Operation at the fundamental will provide more stable output for the upcoming, Navy-funded experiments by the Aerospace Corp.  As noted last week, the installation of these mirrors allows us the opportunity to push the 1.6 and 1.06 micron mirrors to their limits after the THz traps were installed. A presentation to receive JTO funding for the second generation deformable mirror design was prepared.  It will be given later this month.

We prepared several 1.06 micron mirrors to be shipped to the University of New Mexico for measurements of nonlinear effects in the coatings, especially for mirrors that had been used in the FEL vs a spare that had not.

For the standard water cooled mirrors that were replaced on Monday,  two of the mirrors are brazed and two are pasted.  By Tuesday these mirrors were aligned and the output coupler vacuum can was purged overnight with hot dry nitrogen.  On Wednesday, OC vacuum was slow to recover, likely due to extensive component replacement and having the vacuum can open to humid air for so long.  By Thursday, the FEL was lasing albeit with unwanted mirror vibration.  Things improved when Dave reduced water flow to the OC mirrors and turned OFF a vacuum turbo pump.

The modified EO cell/pulse picker for the advanced drive laser, with rotated output polarizer, was reinstalled.  Aligning the EO cell and verifying proper beam specifications downstream of the amplifier chain and frequency doubler is the last "optics" step toward installing the ADL at the FEL.  On other ADL related fronts, Al Grippo is using epics to communicate with ADL electronics although some commands need more work.  The RF group is constructing a 74.85MHz source for the ADL and we anticipate delivery soon.

Terahertz:

For most of this week, Gwyn and I have been on travel.  Gwyn was presenting at the American Vacuum Society Meeting.  I am up at the NSLS at Brookhaven to perform far infrared spectroscopy measurements on several materials of interest, repeat some of the noise measurements on the NSLS light source with the new signal analyzer system that I have built and coded, and to complete as much preliminary setup as possible for the magneto-optical tests that we hope to schedule here for January.

# FSU Effort

FSU FEL.  We continued discussion to cover the merger region and identified a proposed layout of the machine which is somewhat changed from previous ideas but with advantages We held a review of the Parmela modeling status of the injector and booster region of the in operational flexibility and physics performance.  We will continue our analysis to see if the design holds together under further scrutiny.

More PARMELA simulation were made of the injector of the under design FSU FEL. A corrected RMS transverse emittance value was used in the simulations. Deliverables of this set of simulations are the initial beam condition for the beam line design downstream of the accelerating module of the injector. Another set of simulations was carried out to study the beam propagation through a merger. A copy of the present JLab FEL merger was used in the simulations. This set of simulations shows acceptable, from the point of view of beam compression in the merger, phase energy correlation of the beam at the exit of the accelerating module. We also had a brainstorm session to discuss the possible machine layout and to determine next steps in the lattice design.