# MEMORANDUM

To:		Distribution

From:		George R. Neil/gpw

Subject:	FEL Upgrade Project Weekly Brief  -  February 26-March 2, 2007

Date:		March 2, 2007

Highlights:

This week we continued machine operations delivering beam sustained above kW levels for many hours at 1.6 microns.  These longer term tests are extremely valuable in providing experience for prolonged operations, with simultaneous steady power levels, and positional stability.  We also ran at 935nm for extended shifts for the ONR funded LIPSS experiment. 

Correction - for the record, last week's 10 degree off-crest number should have been 20 degrees.  Thanks to those who noticed!

Management:

Steve Benson, Dave Douglas, Michelle Shinn, Bob Rimmer and Claus Rode participated this week in the FEL Technical Area Working Group (TAWG) meeting at Brookhaven National Lab. 

Steve Benson also visited the Source Development Lab at Brookhaven, and the NISUS wiggler in particular.  He also looked at hardware installations for the DOE nuclear physics ERL project, which will provide electrons for cooling the Relativistic Heavy Ion Collider. 

Gwyn Williams represented Jefferson Lab at a Department of Energy, Basic Energy Sciences Advisory Committee meeting in Bethesda this week, and also gave a colloquium at Hampton University on Thursday. 

Mike Klopf and Gwyn Williams provided Alan Todd with material for the quarterly THz program review which was on Thursday.

Operations:

The first string spent much of the week in Brookhaven at the TAWG meeting so the second string took over and ran the machine this week.  They managed quite well.  As is always the case when some are put in charge, they learned quite a bit as well. 
   Over the weekend we had another AC power fault, which dropped the clean power for the FEL facility.  We are still exploring why this happens.  The personnel on hand over the weekend were unable to bring the machine back into operation due to some faults caused by the power cycle.  We therefore spent part of Monday operations bringing the machine back into a running state.  One injector RF control module had to be replaced and this changes the phases in the injector so we started by re-phasing the injector.  Once this was done we set up for high power operations at 1.6 microns.  This power was used for some metal cutting experiments in User Lab 2 and nanotube studies in User Lab 1.  Once these experiments were done we switched to 940 nm operation and started delivering beam for the LIPSS experiment.   This continued for the remainder of the week.  The machine ran stably all week.  We had no problem delivering 1 kW to both the metal cutting and nanotube studies.  The 900 nm mirror losses prevent us from delivering more than a few hundred watts to the LIPSS experiment but it is quite stable and runs for many hours with only occasional interventions to tweak the optical transport or electron beam energy.

WBS 4 (Injector):

The photocathode gun delivered 153 Coulombs and 46 hours of CW beam time for FEL ops during the week.  The high voltage power supply tank was grouted to the floor of the Gun Test Stand.  Fabrication of the resistor assembly for the gun test stand has started.  Progress continues on the final fabrication stage of the new gun chamber.  Preparations are underway to vacuum fire the gun chamber here in the lab, instead of sending the chamber to an outside vacuum furnace.  D. Bullard continued working on coating the wiggler chamber.

Gun HVPS - It remains fully operational.

WBS 6 (RF):

RF - When the machine experienced a hard AC power crash this past weekend, the RF system lost 1 low level RF Control Module out of 28 units.  A recently calibrated unit was installed, the RF parameters downloaded, the machine was then aligned, and the system was operational again.

The current trip level was increased for the faulting external circuit breaker and all of the major individual circuit breakers were lowered to about mid-capacity or less.  This was done to determine which, if any, of the major power supplies is causing the problem.

By analyzing his data of the RF system, T. Powers was able to determine the Probe cables for cavities 5 & 6 of zone 3 are the likely culprits of minor energy jumps.  These cables will be checked and tightened at the next vault access opportunity.  A linac energy signal was developed and displayed on a strip tool screen, so a minor energy change is now easily correctable by the operator.

WBS 8 (Instrumentation):

This week began recovering from the power outage last weekend. A meeting was held with plant services and the electrical safety folks, here is the details from Todd Kujawa (Electrical Safety): the instantaneous trip setting on FEL-BKR-414 was changed from 1.5 (3,000 AIC) to 10 (20,000 AIC) at the request of the FEL group and to be able to further investigate the suspected nuisance tripping (most recent trip was this past weekend). The breakers at PNL-FEL-ELEX were also set back to a point just above halfway on the instantaneous setting. If the breaker was already below this new set point then it was left with its original setting. The expected scenario would be, (if there was a trip due to a high inrush of current), that the trip would occur closer to the equipment that caused the trip instead of at FEL-BKR-414. The end result would be to have FEL-BKR-414 coordinated correctly with PNL-FEL-ELEX.  This setting change at FEL-BKR-414 now has a direct effect of how the LTT process is carried out with PNL-FEL-ELEX since the hazard energy level, with the new setting at FEL-BKR-414, is beyond a level of PPE to protect against while performing voltage verification. The new process will be as follows for LTT of PNL-FEL-ELEX: 
1) turn FEL-BKR-414 instantaneous setting from 10 to 1.5 
2) lockout/tagout FEL-BKR-414 
3) proceed with normal LTT procedures for PNL-FEL-ELEX while wearing the PPE required for a Cat 1 hazard at an energy level of 2.4 cal/cm2. 
4) After work has been completed and FEL-BKR-414 has been re-energized then change its instantaneous setting from 1.5 back to 10 and verify this setting has been saved. 
    I (Todd) have talked to the JLAB electricians about this and facilites management is also aware of this procedure. I (Todd) will place a notice on PNL-FEL-ELEX as a reminder to this procedure. I (Todd) have also talked to the JLAB electricians about installing a VVU on PNL-BKR-414 when the FEL goes down in a few weeks. The VVU would eliminate the need for this extra procedural step when performing LTT on PNL-FEL-ELEX. - MANY thanks to Todd for this help.
    We have begun setting up a new development information server, specifically, we are setting up a JLab Computer Center Maintained Apache server which will be the future location of the configuration control database and devlore content. The machine name will be  "fel.jlab.org" and will also host a wiki of the same build as the now famous Wikipedia website (namely, the "mediawiki platform). This is motivated by the over-due need to move-away from the FEL maintained Microsoft Server platform (with IIS, ASP and MS Access database back-end) which currently runs the site at: "laser.jlab.org". Over the next few months, the existing devlore web-interface will be made obsolete and the functionality of devlore will be provided through the new FEL "Wiki" and related server applications. 
    The opportunity to test the new AMS/Video software has been scheduled for this coming Tuesday owl shift. Even though that is next week we have been continuing to refine it to ensure its success. Also this week, the operational status of the Laser Personnel Safety Systems EPICS data being archived by devlore will be added to the FEL status page at the top of the log book. The graph will provide the last 8 hours of system status. 
    Activities in User Lab 5 began with a re-organization of equipment in order to make room for M. Kelley's JTO collaboration. Fabrication of the User Lab status panels have been completed and tested.  The panels will be installed as time allows. 
    The user runs have forced the staff to become proficient in the many systems required for FEL operations; this experience is invaluable for cross training and for the staff to better understand current operations and how to improve/streamline ops for the future!

WBS 11 (Optics):

FEL Mirror Development

The majority of this week was occupied with my trip and participation in the FEL TAWG meeting at Brookhaven.  I presented a talk on the path to higher average power FELs that was well-received.  The TAWG was also an excellent opportunity to meet with Al Ogloza (China Lake) and Carmen Menoni (Colo. State U.), members of the Optics Improvement Collaboration.  Al brought some very useful information, obtained with Stanford's photothermal deflection technique, that shows microscopic regions of high absorption on an otherwise ultralow loss background.  Identifying the source of these regions, probably due to some process parameter during coating, and eliminating them will be important in the quest to achieve even higher laser damage thresholds than observed now.

Optics Support

This week we supported three user experiments in both setup and FEL operation. The week started with finding the 1.6um OC mirror’s pitch motor some 0.3 radians off its aligned position. We are not sure what caused this dramatic pitch change but we suspect it was due to the weekend power outage and subsequent partial restore of the IOCs.  Monday after re-aligned the 1.6um cavity, we configured Lab 2 for the laser cutting of stainless steel experiment. We closely supported this experiment by providing a dedicated professional optics group person to assist in sample setup, alignment, stage and shutter control, and in experimental direction during the 16 hour shift.

Cavity/Transport Optics

We have finally had time to complete the drawing for the remaining cavity and transport scatter absorbing shields. The drawings are currently in the machine shop for fabrication. We anticipate machining of the molybdenum 3” OC rings for the High-Q cavity cryo-mirrors to be completed next week. Today we will submit a purchase requisition for eight additional 3” deformable HR mirror assemblies. A micro controller and temperature sensors was ordered for the Optical Beam Position Monitor (OBPM). The controller has multiple PWM/timer capabilities. It has been decided that the OBPM wires will only need to be in the beam path for the duration of a measurement request so a timeout is being included in the software. This should eliminate the overheat problems, but a temperature monitor will be included in the software. Preliminary EPICS control requirements for the OBPM chassis interface have been made to the software group.

Laser Work

We did not get the time to work on the advanced drive laser this week due to other work and the available time in Lab 6. Except for one special custom optical assembly, all the parts needed for the Gun Test Stand drive laser phase compensation system are ion the purchase requisition system. We did an initial check on the Ti:sapphire lasers and are in the process of optimization and characterization them. A preliminary plan was made for setting up the laser system for the EO sampling experiment. Our calculation has provided final parameter for the EO sensors. The setup is ready for measuring the response of FEL to modulation of electron bunches. The Antares drive laser is operating normally but the power started to drop with increased current. We will replace the lamps on the nearest possible chance. We helped our Japanese user in setting up the laser-material interaction test and continued to help ODU with the upcoming experiment in Lab 6.

Terahertz:

Much of this week was dedicated to preparations for our presentation at the March APS meeting next week as well as assisting Alan Todd prepare for the quarterly review of the Advanced Energy Systems THz Systems Development program on which we share a collaboration.  The reviews seemed positive, and we hope that the imaging results that we have had so far will help to garner more support for this program.  We will continue to push the THz imaging study, and this week, we received one of the lenses that we expect will improve the performance of our imaging system.  We have another type of lens on order that is not expected to be completed for another 2-3 weeks.  Once we have both lenses, we hope to schedule another THz imaging run and see what improvements we can make. 

Wade Brock, also took data this week for his study on the THz beam polarization.  We had to make some modifications to the setup that was originally planned, but after that, we were able to perform a full set of measurements to check that our system had sufficient dynamic range. We also checked the beam alignment and tested all of the data acquisition systems.  We have not completed the full set of data yet, but the initial results appear to reveal some interesting spatial dependence on the polarization across the beam profile.  We think this can be understood due to the mixture of radially and linearly polarized emission due to edge radiation and synchrotron radiation respectively. Data collection will continue today. 

Also this week, the vacuum FTIR spectrometer was reconfigured and setup to properly interface and align to all of the vacuum systems, the optical table, and the THz beam path after having been removed for some time.  Once the reference laser was realigned to the interferometer, the spectrometer was brought back to life and we are currently running tests on the system.  One problem that was found was that one of the globar sources used for characterizing and testing the system has failed and will need to be replaced.  We are using a spare globar for now, and can continue working on the spectrometer.

Users:

This week we ran 3 user groups.  1.  Eisuke Minehara from JAEA, Japan, is a world expert on high power laser-metal interactions. "In the Stainless steel machining, we were able to use excellent light beam parameters of 0.5-1mmmrad in alignment mode or low power operation, and good parameters of 2-3 mmmrad or larger in high power mode at 1.6 microns.  Four rods of 6mm diameter, 4 pipes of 10mm diameter and 0.5mm thickness, several plates of 5mm thickness, and a nuclear fuel length dummy were used as targets for cutting and drilling.  4 cubes of 10 mm steel with milling marks were exposed to ablate these machining marks. In the very beginning, a couple of rods could be cut reasonably and sharply and welded again instantly.  The FEL power in User Lab #2 was varied between 200W and 1.1kW in Hutch mode.  All of the samples are to be measured and examined by standard material tests and inspections later."  2.  Mike Smith, NASA-Langley writes "NASA is eagerly analyzing samples acquired this week in a new nanoparticle synthesis chamber. Because of the high average power and stability of the FEL, the researchers were able to quickly move through a matrix of over 30 chamber/laser conditions in a single shift.  Some preliminary results show exceptional promise which we hope to confirm and report on shortly."  3.  LIPSS, the ONR funded LIght Pseudo or Scalar particle Search, completed 19 hours of operations at 935nm with the light polarization parallel to the magnetic field.