## Slide 1: Si PiN diode detector for radiation monitoring in CEBAF accelerator service buildings
Vladimir Popov
RadCon
DOE Accelerator Safety Workshop
August 15-17, 2017
Penetration

## Slide 2: Introduction
Radiation detectors for radiation monitoring in proximity of penetrations in CEBAF accelerator service building have been developed studied and tested.
The detector made with commercial BPW34 PIN diodes primarily produced as a photodiode. However, studies show that diodes applicable for x-ray and gamma detection could be used as a cost effective alternative.
The detector has to operate indoor near technological penetrations (RF and He lines) to provide a radiation safety monitoring near penetrations. In total, CEBAF accelerator service buildings have about 126 penetrations.

## Slide 3
Sensor: 10 PIN diodes, model VBPW34FAS, sensitive area = 0.75cm2 total.
Output: pulse, 5V, 10μs
Background count rate: about 2CPM at 10μrem/hr
Sensitivity: 210CPM/(mrem/hr)
Linear response limit: ~60kHz or 17.5rem/hr
Operation limit: ~65kHz
Power: +5V, (3…7V)
Size: L80xW54xH23 mm
Detector specs

## Slide 4
VBPW34FAS PiN Diode
FEATURES
• Package type: surface mount
• Dimensions (L x W x H in mm): 6.4 x 3.9 x 1.2
• Radiant sensitive area (in mm2): 7.5
• Daylight blocking filter matched with 870nm to 950nm emitters
• Capacitance: ≈ 70pF at Vb=0V,
and ≈ 10pF at Vb=60V
• Max Vb=60V

## Slide 5
Si dielectric constant K=11.7. 8pF capacitance means about 100μ depletion zone size at 60V reverse bias. Average electron-hole production energy W=3.6eV, minimum ionization particle energy loss in silicon 39keV in 100μ of Si, and it will produce about 10800 e- or 1.7fc charge from cosmic radiation.
CSA with 2pF cap. produces ≈ 0.8mV output amp. from cosmic muon.
PIN diode sensitivity consideration

## Slide 6
The PIN diode detector circuit block diagram
+3…6V DC
V reg.
SW, +Vcc
+Vb
+60V
L. Reg
Pos.
L. Reg
Pos.
+5V (digital)
+5V
P/Z, LPF
and amp.
2pF
Vthr ≈ 100…200mV
0.8mV
10 μs, 5V
Discr.
Monost. Mult.
TTL
Output
PW Supply
Over Voltage Protection
Ten BPW34 diodes
MCA test point
OPA140

## Slide 7
Chart: 
109Cd, 88keV
133Ba, 81keV
BGR, cosmic
Pulse Height Distributions (PHD), obtained with gamma sources and background-cosmic radiation
Suggested threshold

## Slide 8
Chart: 
| 200 | 100 | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625
 | 43097.04 | 21384.48 | 10941.84 | 5531.76 | 2811.36 | 1395.36 | 695.76 | 324
 | 41868.72 | 20764.56 | 10764.96 | 5465.76 | 2727.12 | 1364.88 | 723.12 | 342.24
Two detectors test-calibration with 137Cs source at Radcon range
(1msv/hr)

## Slide 9
Chart: 
| 4 | 20 | 100
AVR | 865.8333 | 4323.1667 | 21096.1667
D4 | 870 | 4443 | 21681
D5 | 870 | 4218 | 20971
D6 | 871 | 4439 | 22030
D7 | 834 | 4086 | 19891
D8 | 909 | 4656 | 21877
D9 | 841 | 4097 | 20127
Six detectors comparison

## Slide 10
Chart: Linearity limit
| 1E-3 | 1 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 55 | 60 | 65 | 70 | 75 | 80
Rate vs Time | 1 | 999.8333 | 4990 | 9995 | 14982.5 | 19967.5 | 24988.3333 | 29976.6667 | 34953.3333 | 39948.3333 | 44928.3333 | 49860 | 54897.5 | 57468.8333 | 57386.6667 | 35350.8333 | 36584.1667 | 33343.1667
Linearity and response limit of the detector. Tested with LED light pulse.

## Slide 11
OSH Park Order Details
Detectors cost estimate
PCBs: PIN diodes board, $8.05 per batch of three;
Main board, $33.80 per batch of three;
Set of PCB’s for one detector = $13.95;
1455C801BK, aluminum enclosure = $12.19;
Parts ≈ $10;
Component placement in EESSUP, Jlab,
about 2hr per det ≈ $63.00
Total cost of one prototype detector production ≈ $100;
The cost of one will be lower, in case of production larger series

## Slide 12
THANK YOU!

## Slide 13
VB=60V
CSA
+V
Gnd.
Gnd.
TTL Out.
Thr.
+5V
A
B
PIN diode detector layout
C

## Slide 14
137Cs
Cosmic, bgr
PIN detector, pulses at discriminator input

## Slide 15
Chart: 
BGR
133Ba
137Cs
109Cd
Ch.
Gamma spectra

## Slide 16
George Kharashvili, FLUKA simulation