## Slide 1: SVT Stepper Motor Test
Takashi Maruyama, SLAC
HPS Collaboration Meeting
June 16-18, 2014

## Slide 2: SVT and Linear Stage
Wire Frame
L1
L2
L3
Linear Stage is used to
Move SVT layers 1-3
Do wire scan to find the beam
Positioning error
Backlash
Vacuum effect
Long term stability
Test in magnetic field
As we want to place the layer 1 physical edge at only 500 m from the beam, we aim for less than 50 m of positioning error and reproducibility.
Tests:

## Slide 3: Linear Stage and Stepper Motor
10 teeth / 1”
0.100” per revolution
400 encoder pulses per revolution
6.35 m / pulse
0.1” per turn
1 mil per unit
“Home” is at the out
position.
Encoder
Newport XPS Controller
Hovanes provided the stepper motor
parameters.

## Slide 4: Positioning error measurement using Heidenhain MT1200
Heidenhain MT1200
0.2 m accuracy

## Slide 5
350 m from vacuum
120 m backlash
25 m backlash
0 mm
5
10
No vacuum
Under vacuum
Positioning error is  10 m.
Directional backlash is 120 m w/o vacuum and 25 m w/ vacuum.
Vacuum distorts the stage as much as 350 m.

## Slide 6
W=0
W=3.2 lbs
W=5.4 lbs
W=7.4 lbs
380 m
367 m
310 m
313 m
W=5.4 lbs
W=3.2 lbs
W=0
Heidenhain
Weight
Vacuum Effect

## Slide 7
Vacuum
SVT weight
Bottom Stage
Bellows stretched
Bellows compressed
Larger vacuum effect
Larger vacuum effect
Smaller vacuum effect

## Slide 8: Flange stability
Brown&Sharpe
4.2 lbs weight
Pumped to 2×10-5 Torr and valved off
0.5 mil initial settlement
Stable within 0.5 mil

## Slide 9: Stepper motor in B-field
15” , 24”
18”
90 valve
The stage was pumped down in the lab, and was valved off.

## Slide 10
Distance from magnet | Central Field | Field at Stepper Motor
15” | 0.3 Tesla | 100 Gauss
15” | 0.6 | 200
15” | 1.0 | 310
24” | 1.0 | 120
Repeat stage move 5x: 0  20 mm  0

## Slide 11: Conclusions
Positioning error is 10 m.
Directional backlash is 25 m.
  Always move the stage towards the beam.
Stage is stable within 0.5 mil.
No problem in operating in B-field.
Vacuum causes as much as 500 m of stage movement.
  This is not a problem as we have to find the beam using SVT wire scan.
  But we have to remember this effect when we set the limit switch on the linear stage.