## Slide 1: ECal: New APDs testing and gluing
Work performed at JLab by
Emma, Holly, Kyle, Volodymyr
with the participation and/or input from
Alessandro, Andrea, Annie, Gabriel, M., Raphaël, Stepan & INFN
Michel Garçon HPS collaboration meeting, June 2014

## Slide 2: Phases of work
Old APDs & fiber connectors ungluing
Crystal cleaning + (re)wrapping
Frame cleaning
APDs testing and grouping
APDs gluing
Crystal + APD test
Preparing and gluing LED holders
Mounting crystals in frames
Feb. - March
Feb. – June
April
April – May
April – May
May – June
June
June - July

## Slide 3: APD testing
Dark and Light current (LED off, on)
3 Temperatures: 20°C, 18°C, 16°C
Voltage (from 0 to breakdown)
Test bench built in Genova,
Analysis method developed in Roma
(A. Rizzo et al., Report on (CLAS12) Forward Tagger LAAPD Benchmarking)
Present analysis performed by Holly (cross-check Andrea)

## Slide 4: APD quality control
Idark vs Gain linear at all T.
1/G·dG/dV linear fn of G,
indt of T (~ 6%/V).
Gain = G(αV – βT).

## Slide 5: Comparison with Hamamatsu’s numbers
Vbias_Hamamatsu vs Vbias_measured
A few batches of 12 are off by exactly 1 V compared to Hamamatsu’s delivery sheets.
We have no knowledge of the uncertainty in their measurements,
but ours repeatedly gave the same values (within 0.2 V).
Knowing our uncertainty,
we grouped using the voltages and gains that we measured.
VMeasured

## Slide 6: APD selection and grouping criteria
Fit Gain = G(αV – βT)
→ each APD characterized
by the ratio α/β (~ 1.25).
APDs are then grouped by minimizing the spread of voltages V(G=150, T=18°C) of the resulting group

## Slide 7: APD grouping
Gain of the 442 APDs,
each at its group average voltage:
Spread of gain (max-min)
in the 52 groups:

## Slide 8: APD grouping
Grouping based on voltage measured for each APD at Gain = 150, T=18°C.
Groups below show the assigned group number, average voltage for the group, and gain variation of the group at this voltage:

## Slide 9: APD gluing
Cleaning crystal
Pressing crystal onto the APD
Putting glue on the APD
Glue curing overnight
(batch of 20)

## Slide 10: Crystal + APD post-gluing test
Crystal in thermalized box with light injection at the bottom
and preamp on top
Box closed (18°C)
Signal amplitude measurement
at set group voltage

## Slide 11: Glue check
Histogram of
(APDtesting - APD180)/APD180
Limitations:
HV set within +/- 0.5 V, due to integer read-out.
1V difference on HV corresponds to ~0.06 variation this plot.

## Slide 12: LED holders

## Slide 13: Mounting crystals in frames

## Slide 14: Current status
APDs testing – completed (516 tested, 442 selected, 1 broken)
APD grouping for motherboards – completed
Glue new APDs to crystals – completed
Glue check – completed
All crystal wrappings inspected and re-wrapped if needed.
Left to do (this week ??):
  Punch holes in VM2002 on LED holders – started, ~ 15% done
  Glue LED holders to crystals