## Slide 1
Commissioning of a Nebraska-type Retarding Potential Mott Polarimeter
J. McCarter, M. L. Stutzman, T. J. Gay, K. Trantham, P. Adderley, J. Brittian, J. Clark, A. Cook, J. Grames, J. Hansknecht, M. Poelker
Discussion
Introduction
Compact retarding field Mott polarimeters are widely used to measure electron beam polarizations without the high beam energies and radiation hazards associated with conventional Mott polarimeters. The Micro-Mott Polarimeter designed at the University of Nebraska -Lincoln operates at voltages up to 30 keV and has been designed for high efficiency and simplicity of machining and assembly. This polarimeter will be used to qualify photocathode material for the CEBAF machine and for photocathode research. Preliminary results from the commissioning of this polarimeter are presented, including measurements of asymmetry and polarimeter efficiency. Efforts to determine the Sherman function by using electron beams of known polarity from three different photocathode materials are also presented.
Results
Effective Sherman Function
    Determined by comparing measured Mott asymmetries to the known polarizations, as measured in CEBAF, of electron beams from three different GaAs photocathode materials
Efficiency
    Defined as the ratio of the current on the detectors to the total current hitting the target
    Bias of 300 V applied to target to measure current on target, then HV applied and counts on detectors measured. Current detected did not increase with biases from 300V to 7kV, indicating that the efficiency measured up to 30kV is accurate
    Incident current kept low enough to avoid nonlinear dead time issues
Commissioning Process
Measured CEBAF polarizations
33% 77.5 % at 850 nm 80%
43 % at 773 nm
Isometric cut away view of the cathode and electrostatic lens in the sample chamber (SIMION)
Polarimeter commissioning nearing completion
  Electronic setup robust in that different DAQ line configurations give same experimental asymmetry
  Beam steering effects can diminish polarization; unscattered electrons can overwhelm the signal if certain criteria are not followed
  Measured asymmetries and “known” polarization of electron beams coming together toward coherent Sherman function vs. energy
Efficiency and figure of merit (F.O.M.), are larger at low retarding fields. The increase in measured asymmetry does not offset the massive decrease in electrons detected. F.O.M. peaks around 15keV for the same reason. This implies even inelastically scattered electrons still have polarization information.
Average Sherman function of .23 at 30keV is comparable to that measured with gold targets in other designs. The current efficiency and F.O.M. values calculations are roughly a tenth of those reported in the literature.
We show the dependence of F.O.M. on the retarding field for the first time.
We offer a unique and more accurate determination of Seff, using our well characterized electron beam polarizations from different photocathodes
Apparatus
Polarized source with load locked system for cathode exchange uses standard CEBAF heating & activation.
Electrostatic lenses in the source chamber bend beam 90o so it strikes the target with transverse polarization
Laser supplies circularly polarized light, while insertable λ/4 plate changes helicity to decrease systematic errors.
Aperture separating source and polarimeter chambers helps collimate beam into electrostatic lenses.
Two concentric hemispheres accelerate electrons to 30keV and decelerate to incident energy after scattering from gold target.
Retarding field grids in front of Channel Electron Multipliers (CEMs) can repel all but elastically scattered electrons.
Output of CEMs capacitively coupled into
data acquisition system (DAQ).
DAQ: Preamp to discriminator to scalars with LabView to control and count
Mott target and lens assembly
Target
CEMs
Electrostatic lenses
Retarding field grids
View of the entire lens system. Source is
on the left, target on the right. (SIMION)
Chart: Retarding Potential Field (V)
| -0.52600000000000002 | 8.713000000000001 | 18.988999999999983 | 29.141999999999999 | 39.297000000000011 | 49.446000000000005 | 59.639000000000003 | 69.718999999999994 | 79.983000000000004 | 90.10599999999998 | 100.27 | 110.505 | 120.661 | 130.75399999999999 | 141.00900000000001 | 151.14299999999997 | 161.28100000000001 | 171.446 | 181.60499999999999 | 191.82900000000001 | 201.99700000000001 | 212.15300000000002 | 222.29499999999999 | 232.45800000000008 | 242.7 | 252.85700000000008 | 263.05900000000008
 | 0.19788 | 0.20126 | 0.20422 | 0.20535 | 0.20755 | 0.20764 | 0.2092 | 0.20926 | 0.21221 | 0.21157 | 0.2121 | 0.21396 | 0.21495 | 0.2167 | 0.21618 | 0.21978 | 0.22108 | 0.22356 | 0.22297 | 0.22885 | 0.23115 | 0.23114 | 0.23481 | 0.23603 | 0.24013 | 0.25001 | 0.25635
Sample of the effective Sherman function vs retarding potential. Data taken at 30keV incident beam energy from a superlattice cathode.
Seff vs incident beam energy for the four electrons beams of different polarization measured is above. Error bars represent statistical and systematic. Asymmetry vs energy for the same samples is below.
Chart: Incident Target Energy (keV)
| 5 | 10 | 15 | 20 | 25 | 30
 | 0.0015031 | 0.0010645 | 0.00076458 | 0.00057836 | 0.00045398 | 0.00034758
Maximuim efficiency vs the incident beam energy. Error bars reflect errors in measuring incident current and statistics.
Chart: Incident Target Energy (keV)
| 5 | 10 | 15 | 20 | 25 | 30
Bulk | 0.06 | 0.12515 | 0.15485 | 0.17955 | 0.19288 | 0.21015
Superlattice | 0.066 | 0.1495 | 0.18213 | 0.20375 | 0.21981 | 0.23963
Strained at 773nm | 0.08715 | 0.15514 | 0.17897 | 0.20748 | 0.23598 | 0.24065
Strained at 846nm | 0.068065 | 0.15116 | 0.18606 | 0.20935 | 0.22839 | 0.24439
Chart: Incident Target Energy (keV)
| 5 | 10 | 15 | 20 | 25 | 30
Bulk | 1.98 | 4.13 | 5.11 | 5.925 | 6.365 | 6.935
Superlattice | 5.28 | 11.96 | 14.57 | 16.3 | 17.585 | 19.17
Strained at 773nm | 3.73 | 6.64 | 7.66 | 8.88 | 10.1 | 10.3
Strained at 846nm | 5.275 | 11.715 | 14.42 | 16.225 | 17.7 | 18.94
Chart: Incident Target Energy (keV)
| 5 | 10 | 15 | 20 | 25 | 30
Sherm and Eff at 0 retarding | 5.9703e-08 | 1.8045e-07 | 1.8926e-07 | 1.8598e-07 | 1.754e-07 | 1.5256e-07
Max Sherm and Max Eff | 7.4291e-06 | 2.2454e-05 | 2.355e-05 | 2.3142e-05 | 2.1826e-05 | 1.8984e-05
Figure of Merit (F.O.M.) vs incident beam energy. Red (top) data set was taken with the maximum values of the Sherman function and efficiency. The blue (bottom) set was obtained from the values measured at 250 V retarding field.
Superlattice of
GaAs and GaAsP
Strained Layer
GaAs on GaAsP
“Bulk” GaAs
No strain structure
Chart: Retarding Potential Field (V)
| -0.52600000000000002 | 8.713000000000001 | 18.988999999999983 | 29.141999999999999 | 39.297000000000011 | 49.446000000000005 | 59.639000000000003 | 69.718999999999994 | 79.983000000000004 | 90.10599999999998 | 100.27 | 110.505 | 120.661 | 130.75399999999999 | 141.00900000000001 | 151.14299999999997 | 161.28100000000001 | 171.446 | 181.60499999999999 | 191.82900000000001 | 201.99700000000001 | 212.15300000000002 | 222.29499999999999 | 232.45800000000008 | 242.7 | 252.85700000000008 | 263.05900000000008
Efficiency | 0.00076458 | 0.0007101 | 0.00059343 | 0.00052086 | 0.0004641 | 0.00040338 | 0.00036956 | 0.00031513 | 0.00028003 | 0.00025873 | 0.00022622 | 0.00020497 | 0.00017538 | 0.0001593 | 0.00014297 | 0.00012748 | 0.00011003 | 9.5538e-05 | 8.1786e-05 | 6.7559e-05 | 5.637e-05 | 4.7087e-05 | 3.7646e-05 | 2.9e-05 | 2.153e-05 | 1.4067e-05 | 5.9001e-06
h | 1.1744e-05 | 1.1595e-05 | 1.0437e-05 | 9.6045e-06 | 8.7757e-06 | 7.8641e-06 | 7.4234e-06 | 6.4435e-06 | 5.8421e-06 | 5.5317e-06 | 4.9198e-06 | 4.6032e-06 | 4.0464e-06 | 3.8123e-06 | 3.5383e-06 | 3.1908e-06 | 2.8638e-06 | 2.5365e-06 | 2.2477e-06 | 1.9468e-06 | 1.6752e-06 | 1.4298e-06 | 1.2033e-06 | 9.7922e-07 | 7.3313e-07 | 5.0088e-07 | 2.2749e-07
Efficiency (defined as I/I0 ) and the figure of merit, h, (defined as I/I0*Seff2) vs the retarding potential. Data taken at 15 keV from a superlattice cathode.
Asymmetry vs incident beam energy for the same samples as above, with the same legend as above.