## Slide 1: Development of superconducting sensor arrays at NIST
Joel Ullom
Group Leader, Quantum Sensors
NIST
joel.ullom@nist.gov

## Slide 2
Why care? Diverse applications
single optical and near-IR photon sensors for quantum information
millimeter and submillimeter bolometers for astronomical imaging
x-ray sensors for materials analysis and astronomy
γ-ray sensors for nuclear materials analysis
also: dark matter searches, electrons, alpha particles, neutrino mass, …
calorimeters
bolometers

## Slide 3
Review
Low temperatures (typically sub-Kelvin) suppress thermal noise, allowing performance of thermal sensors to surpass more common semiconducting sensors
Not all low temperatures sensors use superconductivity, but most do
Superconductors are extremely useful in sensor circuits:
  as thermometers
  as small gap sensor materials (D ~ 1 meV vs 1 eV for Si)
  for SQUID ammeters
  to make interconnections with high kelectrical/kthermal
  for low specific heats
  …
High Tc superconductors are unlikely to be helpful

## Slide 4
Transition-Edge Sensor (TES) bolometer
0.06
96
96.2
temperature (mK)
resistance (Ω)
0.00
95.8
thin-film thermometer
TES micrograph
SiN
current
P

## Slide 5
Transition-Edge Sensor (TES) calorimeter
0.06
96
96.2
temperature (mK)
resistance (Ω)
0.00
95.8
thin-film thermometer
TES micrograph
SiN
current

## Slide 6
Why care?
Superconducting sensors can provide a unique combination of energy sensitivity and efficiency
240Pu
239Pu
energy-dispersive
gamma-ray
detectors
TES

## Slide 7
Compare to existing EDS technologies
Performance metrics for TES calorimeters are attractive
NIST-Los Alamos gamma-ray spectrometer
256 sensors, time-division SQUID multiplexing
5.76 cm2 active area, ~35% efficiency at 100 keV
Performance limits
from stat mech
 | sup. | semi.
6 keV (x-ray) | 1.1 eV | 120 eV
100 keV (g-ray) | 4.6 eV | 400 eV
5 MeV (a) | 33 eV | 8 keV
detector types
Esat
calculated limits are far better than Si or Ge

## Slide 8
NIST-Los Alamos gamma-ray spectrometer
256 sensors, time-division SQUID multiplexing
5.76 cm2 active area, ~35% efficiency at 100 keV
Performance limits
from stat mech
 | sup. | semi.
6 keV (x-ray) | 1.6 eV | 120 eV
100 keV (g-ray) | 22 eV | 400 eV
5 MeV (a) | 700 eV | 8 keV
detector types
Esat
achieved values still much better than Si or Ge
Compare to existing EDS technologies
Performance metrics for TES calorimeters are attractive

## Slide 9
What makes up a superconducting spectrometer?
1 m
sensors and collimator
readout circuitry
cryogen-free 50 mK platform

## Slide 10
What makes up a superconducting spectrometer?
1 m
readout circuitry
cryogen-free 50 mK platform
CLK DFB RA16 QDC power
digital feedback electronics
and software, lots and lots of software …
sensors and collimator
Technology overview:
Bennett and Ullom, Supercond. Sci. Technol. 28 (2015) 084003

## Slide 11
Simple, robust cryogenics enable dissemination of cryogenic sensors
Designed & commercialized pulse tube-cooled adiabatic demagnetization refrigerator compatible with up to 256 pixels
Several 10’s of units sold worldwide
Taking steps towards next-generation refrigerators compatible with kilopixel arrays
Recent: no liquid cryogens to reach 100 mK
1 m

## Slide 12
Core technology: SQUID readout
SQUIDS = versatile, low noise current amplifiers that provide useful signal gain between superconducting sensors and conventional electronics
0.2 mm
SQUID current amplifier
V
I
T, R, and I change
TES bias circuit
NIST SQUIDS fabricated into superconducting integrated circuits for readout of sensor arrays
32:1 and 1,312:32 multiplexer circuits for TDM or Φ-CDM
75 mm
20 mm
(Stiehl, APL 2012)
SQUIDs enable multiplexed readout: combining the signals from multiple sensors into a smaller number of amplifier chains
  time-division multiplexing
  code-division multiplexing
  frequency-division multiplexing

## Slide 13
Microwave Kinetic Inductance Detectors (MKIDs)
UCSB 10,000 sensor optical array (R ~ 10)
￼
pair breaking in superconducting microresonator changes its microwave properties
a quarter-wave resonator
figs from P. Day, Nature, 2003
naturally compatible with GHz frequency division multiplexing
NIST dual-polarization submm MKIDs for BLAST-TNG balloon mission
an alternative sensor with a distinctive readout
pixel layout
test setup

## Slide 14
Applications - too many to cover all
Polarization in the CMB
Submillimeter astronomical imaging
Dark matter detection
Submillimeter terrestrial imaging
Single near-optical photon counting for quantum information
Synchrotron X-ray science
Industrial X-ray materials analysis
Time-resolved X-ray absorption spectroscopy
X-ray astronomy
Nuclear materials analysis
Neutrino mass measurement
NIST is active in all these areas

## Slide 15
Application area: transforming cosmology into a precision science
angular power spectrum of Cosmic Microwave Background (CMB)
Polarization in the CMB
Submillimeter astronomical imaging
Dark matter detection
Submillimeter terrestrial imaging
Single near-optical photon counting for quantum information
Synchrotron X-ray science
Industrial X-ray materials analysis
Time-resolved X-ray absorption spectroscopy
X-ray astronomy
Nuclear materials analysis
Neutrino mass measurement
Polarization in the CMB
IF detected, primordial B modes are a signature of gravity waves and an inflationary epoch in the early universe

## Slide 16
Atacama Cosmology
Telescope Polarimeter
(1of 3 arrays)
South Pole Telescope Polarimeter
NIST plays an important role

## Slide 17
corrugation profile
Hubmayr et al. JLTP 2012
NIST technology: micromachined feedhorn arrays

## Slide 18
McMahon et al. JLTP (2011)
Four TES per spatial pixel: two frequency bands, X/Y polarization per band
7mm
NIST technology: multichroic, polarization-sensitive detector arrays

## Slide 19
Millimeter-wave polarimetry: status and prospects
NIST parts in almost all instruments in the field
high profile science
first deployed multichroic array
currently working on large new instrument: AdvancedACT
  migration to 150 mm wafers
  lower loss dielectrics
  more frequency bands
also, starting work on SPIDER balloon payload
also, starting work on planning for LITEbird satellite
future of field is bright
press after announcement of BICEP2 results, but since shown NOT to be primordial B modes

## Slide 20
Application example: synchrotron science
Polarization in the CMB
Submillimeter astronomical imaging
Dark matter detection
Submillimeter terrestrial imaging
Single near-optical photon counting for quantum information
Synchrotron X-ray science
Industrial X-ray materials analysis
Time-resolved X-ray absorption spectroscopy
X-ray astronomy
Nuclear materials analysis
Neutrino mass measurement
Synchrotron X-ray science
there are many different synchrotron measurements. Superconducting sensors are useless for most.
but, there are applications in photon-in, photon-out spectroscopy when outgoing photons are scarce
X-ray emission spectroscopy (XES): energy of fluoresced photons provides elemental, structural, and electronic information. The last two require very good (eV-scale) energy resolution.
high collecting efficiency lets you look at trace samples, radiation-sensitive samples, and dynamic effects
also, absorption spectroscopy and x-ray scattering
NIST x-ray microcalorimeter spectra of two explosive compounds

## Slide 21
Material science at large facilities
NIST spectrometer at the NSLS
NIST spectrometer at the APS
Ullom et al cover article in Aug. 2014 Synchrotron Radiation News
also, system under construction for SSRL, delivery in ~ January 2016

## Slide 22
Polarization in the CMB
Submillimeter astronomical imaging
Dark matter detection
Submillimeter terrestrial imaging
Single near-optical photon counting for quantum information
Synchrotron X-ray science
Industrial X-ray materials analysis
Time-resolved X-ray absorption spectroscopy
X-ray astronomy
Nuclear materials analysis
Neutrino mass measurement
Application example – industrial X-ray analysis
Electron beam-induced X-ray fluorescence is a ubiquitous tool for determining elemental composition on SEMs and TEMs
Silicon drift detectors (SDDs) provide high rates, large area, but ~ 120 eV FWHM resolution
Wavelength dispersive detectors are evolving. Modern units provide high rates, medium effective area, but complicated efficiency curve and ~ 10 eV resolution
TES arrays can provide medium rates, small effective area, but simple efficiency curve and < 3 eV resolution
Cryogen-free
spectrometer with 16 TES X-ray array –
STAR Cryoelectronics
Industrial X-ray materials analysis

## Slide 23
Chemical shift detection with commercial spectrometer
Polarization in the CMB
Submillimeter astronomical imaging
Dark matter detection
Submillimeter terrestrial imaging
Single near-optical photon counting for quantum information
Synchrotron X-ray science
Industrial X-ray materials analysis
Time-resolved X-ray absorption spectroscopy
X-ray astronomy
Nuclear materials analysis
Neutrino mass measurement
Industrial X-ray materials analysis
Los Alamos/STAR Cryo
uranium chemical shifts on commercial SEM-TES instrument
several STAR spectrometers sold and installed at customer facilities

## Slide 24
Looking to the future: how to reach arrays of 106 elements?
1 mm
n1, n2, n3
1. superconducting resonators as detectors:
Microwave Kinetic Inductance Detectors (MKIDs)
2. resonators with RF-SQUIDS and TESs
stimulus changes power at n1 that reaches amplifier
shared feedline
resonators
RF-SQUIDs
input flux shifts GHz resonance frequency
Single microwave amplifiers can provide GHz of readout bandwidth
Likely path for next-generation instruments

## Slide 25
18,900 sq. ft. state-of-the-art class 100 clean room
Opened in 2012
$20-$25M of tools, $30M of infrastructure
Precision Imaging Facility:
Aberration corrected TEM
Dual beam focused ion beam mill/scanning electron microscope
Helium ion microscope
Local electrode atom probe
Boulder Micro-fabrication Facility:
60% of site research passes through
18,900 sqft class 100 cleanroom
Full lithography suite (ASML 0.35 µm stepper, reticle pattern generator)
GaN and GaAs two chamber MBE
15 deposition tools
13 dry etch tools
8 tube furnace bank
10 nm resolution EBL (future)
NIST Boulder Microfabrication Facility

## Slide 26
Conclusions
Cryogenic detectors provide a unique combination of energy sensitivity and collecting ability.
There are already applications where cryogenic detectors are “required” and anticipated performance improvements will make new applications accessible in the future.
The tools to fabricate and cool cryogenic sensors are well established but their cost and complexity should be acknowledged.
The development of microwave readout techniques is an active area of effort. These techniques are a likely path to arrays of 106 sensors. Today’s arrays contain 102-104 sensors.
NIST is active in most areas of cryogenic sensor technology. But, we are also interested in new projects and new collaborators.
Thanks!
joel.ullom@nist.gov