## Slide 1
FTOF and CTOF Hardware, Calibration, Reconstruction, Simulation Software Status
Daniel S. Carman
Jefferson Laboratory

## Slide 2
FTOF KPP Hardware Status
MON12 System Monitor
Issue with missing TDCs in S3 for KPP run due to missing clock signal
Raffaella De Vita

## Slide 3
FTOF KPP Hardware Status

## Slide 4
Hardware Monitoring Tool
Mode1 pulse display
  ADC / TDC display
    FADC pedestals
FTOFMON
Cole Smith

## Slide 5
CTOF KPP Hardware Status
ADC
TDC

## Slide 6
FTOF Calibration Sequence
HV/Gain Calibration |  | 
Step # | Action | Requirements
1 | Gain Calibration | EC Pixel Trigger
Timing Calibration |  | 
1 | PMT Timing Diff. | Raw TDC data
2 | Attenuation Length | Good Dt
3 | Effective Velocity | Good TBT + Dt
4 | Time-Walk | Good TBT + Dt
5 | P2P | Good TBT + Dt
6 | RF Offset | Good TBT + Dt
Nominal assumption is that for each counter both PMTs have good ADC and TDC values
– code will ultimately be able to complete calibrations for a subset of hardware failures

## Slide 7
Calibration Suites
FTOF Calibration Suite
Louise Clark, Haiyun Lu
CTOF Calibration Suite
java-based, modular suites based on CLAS12 common tools APIs

## Slide 8
FTOF Gain Calibration
panel-1b
cosmic data; KPP data
For each counter fit ADC geometric mean and ADC log ratio distributions
use to determine HV settings to gain balance the two PMTs and to position the minimum ionizing particle peak in a specific ADC channel
use for the energy loss calibration in the reconstruction code

## Slide 9
FTOF PMT Timing Difference
panel-1b
cosmic data; KPP data
Center counter PMT time difference distributions about zero to account for differences in PMT transit times and signal propagation times

## Slide 10
Attenuation Length
panel-1b
cosmic data; KPP data
Use correlation of coordinate and ADC information to determine attenuation length of light for each counter

## Slide 11
Effective Velocity
coorDC = 0.5 (tL – tR) veff
Match hit coordinate from FTOF to hit coordinate from DC tracking to find veff, the effective speed of light propagation in the scintillation bar
panel-1a
panel-1b
KPP data

## Slide 12
Time-Walk Correction
Time-walk functional uses RF time as the reference:
tres = tFTOF – L/b – tRF
Fit tres vs. ADCL,R with function:
l/ADCm
The FTOF system uses leading-edge discriminators for the readout
t=f(ADC) as pulses of different amplitude cross LED threshold at different times
(“time-walk”)
KPP data
l
m

## Slide 13
Time-Walk Correction
An alternative approach to time-walk corrections is to determine calibration parameters that do not depend on the RF or on DC tracking information
Compare the TDC time to the time derived from fitting the leading edge of the flash ADC pulse
Work is underway to compare the two time-walk algorithms and to better understand the systematics
The FADC timing resolution still needs to be improved to be able to employ this approach for corrections
Determine time-walk parameters from the residual:
tres = tTDC – tFADC
Fit tres vs. ADCL,R with function:
l/ADCm
KPP data
Raffaella De Vita, Andrey Kim

## Slide 14
Paddle-to-Paddle Correction
Determine relative time offsets among paddles:
Use exclusive e-p+/- concidences to link times of all counters across the system
Dtv = tev – tpv
P2P = max bin # Dtv * 2.004 ns
t2FTOF = t1FTOF - P2P
Determine initial RF offset for the FTOF system:
For each FTOF paddle compute:
This leaves 2.004 ns ambiguity as beam bucket associated with the event is unknown
t1FTOF = t0FTOF + <DRF>
DRF = tRF – (teFTOF – Le/b)
<DRF> = mod(DRF, 2.004 ns)
Select electrons or pions with b~1
e
p
Algorithm now being developed/tested
Raffaella De Vita, Ralf Gothe

## Slide 15
Paddle-to-Paddle Correction
RF Offset
P2P Offset
Coincidence statistics

## Slide 16
Reconstruction Software
TOF reconstruction code determines:
hit times (tL, tR, <t>)
hit coordinates (x)
deposited energies (EL, ER, <Edep>)
associated time, coordinate, and energy uncertainties
performs hit clustering and matching
combines hit times from panel-1a and panel-1b
- Code designed to function for all “allowable” hardware conditions
- Most of the code validated in detail; work remains on combining hit times
Veronique Ziegler

## Slide 17
Optimizing System Time Resolution
Combine the measured times from the FTOF panel-1a and panel-1b to optimize the resolution
TDR
(20% gain) | e | p | p+ | p-
FTOF 1A | 152 ps | 164 ps | 160 ps | 151 ps
FTOF 1B | 62 ps | 67 ps | 65 ps | 61 ps
FTOF 1A+1B | 49 ps | 54 ps | 54 ps | 51 ps
GEMC Studies with cluster=1
Algorithms for cluster > 1:
use hit with tmin
use hit with Emax
use weighted average
Evgeny Golovach

## Slide 18
Simulation Code
The FTOF and CTOF system modeling in the CLAS12 GEMC simulation code is designed to match the actual hardware response:
ADC digitization:
Proper model for ADC pulse based on realistic modeling of photoelectron statistics
Includes attenuation length effects
Ultimately will use pulse shape to determine time from ADC data
TDC digitization:
Proper model for time-walk effects
Realistic PMT time resolutions
Realistic effective velocity
Include PMT time and P2P offsets
FTOF & CTOF simulations studied in detail during December CALCOM Calibration Challenge
hit-process routines

## Slide 19
Remaining Work
Finish algorithm development for T-W, P2P, and RF offsets
  Test algorithms on GEMC data and KPP beam data
Complete full calibrations of KPP data
  Put constants in CCDB
Continue development of TOF calibration suites
  Optimize calibration suites to improve user experience
  Fitting of 2D histograms not done properly
  Develop suite to calibrate around hardware failures
  Explore alternative algorithms where improvements are promising
Update algorithm documents and prepare calibration tutorials
Add system survey offsets (Dx,Dy,Dz,qx,qy,qz) into geometry service
Hardware work:
  CTOF detector installation, repairs, cabling, cosmic calibrations
  FTOF detector repairs (where accessible)

## Slide 20
Preliminary FTOF Survey

## Slide 21
Summary and Conclusions
The FTOF/CTOF calibration suites are becoming mature codes:
  The calibration suites have been used to calibrate the KPP beam data (gain balance, PMT time diff., attenuation length, eff. vel., time walk)
  Work ongoing to finish development of P2P and RF time offsets
The FTOF/CTOF reconstruction code has been validated through detailed systematic studies of Monte Carlo and KPP beam data:
  All aspects of code functionality studied
  Some minor work remains on combining hit times for FTOF
The FTOF/CTOF GEMC hit process digitization routines have been developed to match the actual hardware response:
  Extensive testing during recent Calibration Challenge exercise
  Work remains on using pulses for FADC modeling
FTOF hardware in good shape (<1% of channels have issues); Significant effort upcoming to install CTOF in solenoid and test. All installation procedures tested during KPP installation