## Slide 1: C75 System Specifications
C75 Implementation and Project Review
15th February 2018, JLab
Frank Marhauser

## Slide 2: Outline
C75 energy gain and contingency and corresponding gradient specifications
Unloaded Q (Q0) specification (prospect of large grain niobium material)
Microphonics allowance with respect to usable generator power
Power input coupler’s (FPC) external Q (Qext) specification
Field emission specification rationale

## Slide 3: Energy Gain and Accelerating Field (Eacc)
Energy gain baseline requirement for C75 cryomodules:
  Accumulated energy gain shall be 75 MeV (when leaving CMTF after acceptance testing)
With 7.5% contingency (similar to C100):
  RF system shall allow operating up to 20.5 MV/m per cavity (considering usable generator power and microphonics)
  Cryomodule may achieve 80.625 MeV
Parameter | Unit | Baseline | with 7.5% Contingency
Energy Gain | MeV | 75 | 80.625
Effective accelerating voltage per cavity | MV | 9.375 | 10.078
Corresponding accelerating field per cavity (on crest acceleration) | MV/m* | 19.07 | 20.5
* Nominal active cavity length: 0.4916 m
Notes: Red Interior Page Design 1

## Slide 4: Q0-Specification (at T = 2.07 K)
Cavity Type | Q0 specification | Eacc (MV/m)
C20/C25 | 2.4e9 | 5
C50 | 6.8e9 | 12.5
C100 | 7.2e9 | 19.2
C75 | 8e9 | 19.07
Prospects to benefit from (medium-purity) large-grain (LG) ingot Niobium
Distribution of quality factors at 2 K just below the quench field limit (over 30 single- and multi-cell cavities) as a function of quench fields for cavities made from CBMM ingot material of different purity. The findings for LG cavities with OC shape are highlighted (red circles) and exceed the specification (star symbol) for the C75 cavity.
Bpk = 80 mT (~86 mT at 20.5 MV/m)
Notes: Red Interior Page Design 1

## Slide 5: C50 Q0-Degradation Experience
What we know so far (C75 prototype cavity pair in C50-13)
C75 prototype cavities show highest Q0-values among all refurbished C50 cavities (close to 8e9)
Notes: Red Interior Page Design 1

## Slide 6: Potential Q0 and Eacc degradation
Cav. # | Cav. Type | Cav. SN | Eacc,max VTA MV/m | Eacc,max CEBAF commiss. MV/m | Q0 VTA | Q0 (2.07K) CEBAF commiss. | Q0 degradation VTA  CEBAF % | Performance limit
1 | C75 HC LG | 5C75-001 | 19.4 | 19.1 | 9.0e9 | 7.6e9 | 15.1 | Quench
2 | C75 HC LG | 5C75-003 | 13.7 | 14.2 | 8.3e9 | 7.7e9 | 7.6 | Quench
3 | C50 OC FG | ia274 | 18.3 | 16.6 | 9.6e9 | 6.5e9 | 32.3 | Quench
4 | C50 OC FG | ia345 | 19.9 | 17.4 | 8.5e9 | 4.3e9 | 49.4 | Waveguide vacuum
5 | C50 OC FG | ia366 | 9.2 | 9.2 | 7.6e9 | 7.0e9 | 25.6 | Quench
6 | C50 OC FG | ia351 | 15.1 | 14.0 | 7.8e9 | 5.8e9 | 25.6 | Quench
7 | C50 OC FG | ia038 | 18.0 | 16.9 | 7.0e9 | 6.0e9 | 14.3 | Quench
8 | C50 OC FG | ia260 | 17.0 | 15.5 | 6.7e9 | 4.5e9 | 32.8 | Quench
Avg. |  |  | 16.3 | 15.4 | 8.1e9 | 6.2e9 | 22.3 | 
VTA and C50-13 commissioning results in CEBAF (1L13) including C75 cavity pair (Mike Drury, Nov. 2017)
For the C75 prototype pair one must consider series of in-house fabrication issues
  Q0 and quench limit related to Electron Beam welding issues (G. Ciovati et al., Cavity Fabrication Tech Note, May 2017)
Lessons learned to be applied with next C75 cavities
  Procurement of C75 cavities: EB welding at commercial vendor may mitigate performance issues (vendors are liable for mechanical fabrication, QC imposed)
Persistent Q-loss in C50 cavities still not fully understood
  After upgrades to ‘magnetic hygiene’ including cavity magnetic shielding
Notes: Red Interior Page Design 1

## Slide 7: Microphonics
What we know so far
  T. Powers, Microphonics measurements for C50-13/1L-13,10/31/2017)
  Microphonic detuning (6·δfrms) < 30 Hz for both C75 cavities
Cavity Location | Cavity SN | Microphonic detuning δfrms Hz | 6· δfrms Hz | Microphonic peak detuning δfpk Hz
1 | 5C75-001 | 4.7 | 28.2 | 11.9
2 | 5C75-003 | 4.8 | 29.0 | 18.2
3 | ia274 | 3.8 | 22.5 | 14.4
4 | ia345 | 3.6 | 21.5 | 14.8
5 | ia366 | 5.8 | 34.5 | 20.7
6 | ia351 | 3.0 | 17.9 | 7.5
7 | ia038 | 3.9 | 23.4 | 14.5
8 | ia260 | 3.9 | 23.3 | 15.3
Average value |  | 4.2 | 25.0 | 14.7
Microphonic detuning measured for C50-13 cavities in CEBAF (1L13). Each of the values are the maximum values measured among several data sets (3-5) taken for each cavity
Notes: Red Interior Page Design 1

## Slide 8: Maximum Usable Generator Power
Maximal generator power (Pg): 8 kW
  Note: Attenuation in WR650 transmission line is ~0.2dB/100 ft.
  We have 15 meter (~50 ft.) transmission line length  ~0.1 dB loss
  Other losses (e.g. insertion losses of waveguide filter, klystron circulator) to be added
  Sum assumed is 0.6 dB (cf. J.A. Fugitt, T.L. Moore, “CEBAF Superconducting Cavity RF Drive System”, typical waveguide losses of 0.5 dB for 15 m)
Maximum usable generator power: 7 kW
Notes: Red Interior Page Design 1

## Slide 9: Input Coupler Qext Specification
Must consider full beam loading (Ib = 460 µA) and up to Eacc,max = 20.5 MV/m
Usable Pg = 7 kW constraints microphonics allowance to ~31 Hz
Optimum Qext around 2e7 with realistic microphonic detuning levels
Parameters used: R/Q = 525.4 , Eacc = 20.5 MV/m, Lact = 0.4916 m, Ib = 460 µA, and Q0 = 8e9
7 kW threshold
Notes: Red Interior Page Design 1

## Slide 10: Qext Specification vs. Microphonics Allowance
Parameters used: R/Q = 525.4 , Lact = 0.4916 m, Ib = 460 µA, and Q0 = 8e9
Based on microphonics measurements in C50-13 for C75 cavity pair (6·dfrms close to 30 Hz), Eacc = 20.5 MV/m can be supported by RF system at full beam loading (Ib = 460 µA), when usable generator power is 7 kW
At Qext = 2e7 the peak microphonics allowance relaxes at smaller accelerating fields (e.g. 38.2 Hz at 19.07 MV/m)
E.g.: A Qext = 1e7 does not support required accelerating fields for a C75 CM at full beam loading
Qext = 2e7 ± 15 %
Notes: Red Interior Page Design 1

## Slide 11: HOM Impedance Specification
Specifications for HOM Impedance Threshold for Beam Breakup Instability (BBU) follow C100 cavity rationales
BBU Impedance threshold for deflecting HOMs: Rtr = 2e10 /m (stretched goal 1e10 /m , cf. G. Krafft et al., JLAB-TN-09-015)
  Definition: Rtr = R/Q(r)/(k·r2)·Ql
  R/Q = HOM’s characteristic impedance, Ql = loaded Q of HOM, r = radial offset, k = wave number
Numerical Analysis done to survey HOM impedance spectrum
  1) Short wakefield calculation to identify crucial HOMs (spectrum not resolved)
  2) Complex Eigenmode calculations to resolve individual HOMs and calculate R/Qs and loaded Qs (used tetrahedral mesh with dogleg and inner tube adapter)
Notes: Red Interior Page Design 1

## Slide 12: C75 Cavity HOM Impedance Spectrum
TE11 beam tube cutoff
FPC cutoff
HOM waveguide cutoff
BBU threshold (12 GeV)
BBU threshold (stretched goal)
Old CEBAF cavity design: In this regime only FPC works as a
HOM damper  Cannot verify in VTA w/o proper setup
Short (~140m) wakefield calculation
(peaks not resolved)
Eigenmode calculations
These HOM have strong monopole component
not subtracted from dipole impedance  worst case
Notes: Red Interior Page Design 1

## Slide 13: Field Emission (FE)
Cav. # | Cav. Type | Cav. SN | Eacc,max VTA MV/m | Eacc,max CEBAF MV/m | FE onset VTA MV/m | FE onset CEBAF commiss. MV/m
1 | C75 HC LG | 5C75-001 | 19.4 | 19.1 | 17.9 | - after RF conditioning
2 | C75 HC LG | 5C75-003 | 13.7 | 14.2 | 9.9 | 10.9
3 | C50 OC FG | ia274 | 18.3 | 16.6 | 15.6 | -
4 | C50 OC FG | ia345 | 19.9 | 17.4 | 15.0 | 9.4
5 | C50 OC FG | ia366 | 9.2 | 9.2 | 8.5 | -
6 | C50 OC FG | ia351 | 15.1 | 14.0 | 14.1 | -
7 | C50 OC FG | ia038 | 18.0 | 16.9 | - | -
8 | C50 OC FG | ia260 | 17.0 | 15.5 | 10.9 | 7.5
Specification follows experience with LCLS-II:
  Provide field-emission free cavities up to quench field limit in the VTA prior assembly rather than defining an upper field limit for FE onset
  Could imply more effort, be prepared for cavity retreatment
  Risk: Potential re-treatment cycles impact availability of cavities for string assembly (schedule) and implies more costs
  First re-treatment is typically high-pressure rinse (high success rate for LCLS-II)
  Improve procedures for clean cavity string assembly
VTA vs. CEBAF commissioning for C50-13 cavities (M. Drury):
Notes: Red Interior Page Design 1

## Slide 14: C75 RF Specification Table
Parameter | Unit | Value | Comments
Energy gain per CM | MeV | 75 | Max. 80.625 MeV with 7.5% contingency, 8 cavities per CM
Operating RF frequency | MHz | 1497 | with cavity under compression
Operating temperature | K | 2.07 | 29 ± 0.1 Torr nominal helium pressure
Maximum beam current | µA | 460 | 
Microphonic detuning δf (rms) | Hz | 5 | 
Microphonic detuning δf (peak) | Hz | 30 | 
Average accelerating field (Eacc,avg) | MV/m | 19.07 | Ueff = 9.375 MV, Epk = 46.6 MV/m, Bpk = 79.7 mT (δf = 38.6 Hz max . allowable)
Maximum accelerating field (Eacc,max = Eacc,avg +7.5%) | MV/m | 20.5 | Ueff = 10.08 MV, Epk = 50.1 MV/m, Bpk = 85.7 mT (δf = 31 Hz max . allowable)
Maximum beam loading | kW | 4.6 | At Eacc = 20.5 MV/m, 4.3 kW at Eacc = 19.07 MV/m
Remanent magnetic field in CM | mG | 10 | After cryomodule degaussing (absolute field)
Allowable RF window losses | W | ≤ 2 | As measured in special setup in Dewar at 2 K (low power) and extrapolated to 5 kW forward power
Q0 at T= 2.07 K |  | ≥ 8e9 | up to Eacc = 19.07 MV/m (incl. all RF losses)
Cavity wall losses (Pc) | W | 20.9/24.2 | for Eacc = 19.07/20.6 MV/m at Q0 = 8e9
Usable generator power (Pg) | kW | 7 | using 8 kW klystrons assuming 0.6 dB attenuation in transmission line from klystron
Qext FPC |  | 2e7 ± 15 % | adjustment to higher Qext-values possible by WR650 stub tuning
Resonant bandwidth (f0/Ql) | Hz | 75 | 
HOM dipole impedances Rtr | /m | ≤ 2e10 | Rtr = R/Q(r)· Ql_HOM/(k·r2) BBU impedance threshold is 2e10 /m for 12 GeV baseline physics with 460 µA max., stretched goal is 1e10 /m
Notes: Red Interior Page Design 1

## Slide 15
Questions ?

## Slide 16
Backup Slides

## Slide 17: CEBAF Actual Cryomodule Energy Reach
CEBAF RF dashboard status: 2/6/2018
Cryomodule Type | Energy Gain (MEV) Avg. ± STD | % of nominal
C20 | 27.5 ± 3.6 | 137
C50 | 41.9 ± 7.3 | 84
C100 | 79.9 ± 3.7 | 80
Notes: Red Interior Page Design 1

## Slide 18: CEBAF Actual Cryomodule Energy Reach
Cryomodule Type | Energy Gain (MEV) Avg. ± STD | % of nominal
C20 | 27.5 ± 3.6 | 137
C50 | 41.9 ± 7.3 | 84
C100 | 79.9 ± 3.7 | 80
C75 | ~60 ? | ~80 ?
CEBAF RF dashboard status: 2/6/2018
Notes: Red Interior Page Design 1

## Slide 19: Qext Specification
At Eacc,max = 19.07 MV/m
Parameters used: R/Q = 525.4 , Eacc = 19.07 MV/m, Lact = 0.4916 m, Ib = 460 µA, and Q0 = 8e9
Notes: Red Interior Page Design 1

## Slide 20: Qext Specification vs. Microphonics Allowance
Parameters used: R/Q = 525.4 , Lact = 0.4916 m, Ib = 230 µA, and Q0 = 8e9
Half beam loading (Ib = 230 µA), Pg = 7 kW
Notes: Red Interior Page Design 1

## Slide 21: Optimum Qext with Beam Current and Microphonics
Parameters used: R/Q = 525.4 , Lact = 0.4916 m, Q0 = 8e9
Nominal Eacc = 19.07 MV/m
Efforts to suppress microphonics by design (cf. talk on cavity design)
Notes: Red Interior Page Design 1