## Slide 1
Daniel S. Carman, Mar. 5, 2019

## Slide 2
CLAS12 Run Periods
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Engineering Run:
2.2, 6.4, 10.6 GeV polarized electrons
Liquid-hydrogen target
Dec. 2017 – Jan. 2018
Run Group A:
13 experiments
10.6 GeV polarized electrons
Liquid-hydrogen target
1) Feb. – May 2018
    1.1 PB, 126 mC
2) Sep. - Nov. 2018
    0.6 PB, 99 mC
  ~40% of approved time
Run Group K:
3 experiments
6.6, 8.8 GeV polarized electrons
Liquid-hydrogen target
Nov. – Dec. 2018
    0.35 PB, 45 mC
  ~7% of approved time
Run Group B:
8 experiments
10.6 GeV polarized electrons
Liquid-deuterium target
Feb. – Mar. 2019
    0.28 PB, 36 mC (so far)
  ~30% of approved time

## Slide 3
Initial calibrations for all systems completed for pass-0 cooking
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
CLAS12 Calibration Status
# | Run | Run Group | Torus | Sol | <i> (nA) | Eb (GeV) | Date
1 | 3222 | A (Spring 18) | +100% | -100% | 25 | 10.6 | 2/15/18
2 | 3355 | A (Spring 18) | -100% | -100% | 35 | 10.6 | 2/20/18
3 | 4013 | A (Spring 18) | -100% | -100% | 50 | 106 | 4/18/18
4 | 5036 | A (Fall 18) | -100% | -100% | 45 | 10.6 | 10/11/18
5 | 5664 | A (Fall 18) | +100% | -100% | 50 | 10.6 | 11/26/18
6 | 5700 | K (Winter 18) | +100% | -100% | 30 | 7.5 | 11/30/18
7 | 6164 | B (Spring 19) | -100% | -100% | 35 | 10.6 | 2/10/19
8 | 6233 | B (Spring 19) | -100% | -100% | 35 | 10.6 | 2/25/19

## Slide 4
CLAS12 Calibration Sequence & Dependencies
DC Calibration:
time  distance calibration
  relies on at least crude ST calibration from FTOF (few ns level)
FTOF (+ CTOF Time Matching) Calibration:
energy calibration
  can be done before DC calibration using even crude DC calibration parameters for path length corrections
timing calibration
  calibrate FTOF timing
  defines event ST using electron in ECAL (1st option), positron in ECAL (2nd option), high momentum pion/track in DC/FTOF (3rd option)
  calibrate offset between CTOF FADC and TDC time
CLAS12 Subsystem Calibration:
CND, CTOF, ECAL, FT (Hodo, Cal), HTCC, LTCC, RICH calibrations
  timing calibration employs event ST from FTOF
  energy calibration employs PID from EB
Assumption: CVT + FT-Trk do not require calibration (only good alignment)
recook
recook
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 5
Monitoring Tools
Andrey Kim
Electrons per trigger per sector
Average HTCC nphe per sector
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
https://clas12mon.jlab.org/rga/timeline/

## Slide 6
When to Recalibrate DC?
Time difference between superlayer pairs  avg. velocity:
R1: 362 mm difference ⇒ ~10 ns
  Avg. velocity ~36 mm/ns
R2: 734 mm difference ⇒ ~40 ns
  Avg. velocity ~18 mm/ns
R3: 860 mm difference ⇒ ~40 ns
  Avg. velocity ~21 mm/ns
R2 and R3 are sensitive to 5 ns effects:
5 ns x 20 mm/ns  100 mm
5 ns in middle of cell  DTmax=10 ns
Plan:
Update Tmax when DTmax = ±5-20 ns
Full recalibration when DTmax>20 ns
R1 numbers  ±3-10 ns and >10 ns
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 7
When to Recalibrate FTOF?
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
p/K | 2.8 GeV
K/p | 4.8 GeV
p/p | 5.4 GeV
Nominal PID separation from design specs at 4s level:
Recalibrate FTOF when <dt> worsens to shift separation by Dp = 250 MeV
Average FTOF dt = 80 ps
⇒ Recalibrate FTOF when:
Panel-1a <dt> ≥ 170 ps (nominal avg. = 150 ps)
Panel-1b <dt> ≥ 90 ps (nominal avg. = 80 ps)

## Slide 8
ECAL Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Study ECAL performance using p0 analysis:
Change in p0 mass with time indicates ~5-10% loss of PMT gain in RG-A spring run
To track PMT calibrations need to match cosmic data to beam data
S1
S2
S3
S4
S5
S6

## Slide 9
ECAL Timing Calibration
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Previous algorithm
New algorithm

## Slide 10
ECAL Timing Calibration
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Previous algorithm
New algorithm
PCAL
ECin
ECout
PCAL
ECin
ECout
U V W
U V W

## Slide 11
Forward Tagger Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Calibration & Performance
FT-Cal:
  Resolutions:
    se/E ~ 3.3%
    sMgg ~ 4.4%
  Optimize shower leakage corrections
  Add time walk correction to timing calibration
FT-Hodo:
  Energy and timing calibrations fully implemented and tested
FT-Trk:
  Track reconstruction implemented and being validated
  Tracker alignment work not yet begun
Single crystal timing resolution < 200 ps with energy cut
Hardware status
FT-Cal & FT-Hodo:
  Stable operations (FT-Cal at 5° C, upgraded FEE for FT-Hodo to reduce gain dependence on beam current)
FT-Trk:
  Cables swaps of 2018 Spring and Fall run understood and fixed in software
  2 disconnects and 2 swaps after January installation: need to improve checkout procedure
  Noisy PLC under investigation
Spring18 FT-TRK occupancy after swap corrections

## Slide 12
Forward Detector PID
positives
negatives
π+
K+
p
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
run 4013, 10.6 GeV, T:-100%, S:-100%

## Slide 13
Timing response limited to date:
  need to turn on vz corrections to RF:
    adds 83 ps in quadrature
  need to turn on position-dependent corrections:
    improves Dt by 15%
  resolution smearing due to TDCs (100 ps LSB):
    panel-1a/2 : all counters
    panel-1b: counters #49-62
Combining times for panel-1b/1a:
  work in progress using MC data:
    improves Dt by 15% from using time from panel-1b alone
  study includes optimize hit cluster definitions in each FTOF panel
FTOF Status
10.6 GeV, Run 4013
Tor:-100%, Sol:-100%
tag 5b.7.1
corr off
corr on
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 14
FTOF PID Limitations
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
4s Limit | Current | Spec
p/K | 2.3 GeV | 2.8 GeV
K/p | 3.9 GeV | 4.8 GeV
p/p | 4.6 GeV | 5.5 GeV

## Slide 15
Central Detector PID
run 4013, 10.6 GeV, T:-100%, S:-100%
positives
negatives
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 16
Time resolution limitations (current):
  Lack of RF vertex correction
  Issues with curved end shapes
  Response effects
CTOF scintillation bar
DNP cooking, RF vzcorr
10.6 GeV Torus=-100%, Solenoid=-100%
Run 4013, tag 5b.4.0
CTOF Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Time "residual" used for calibration:
z = ½veff (tU-tD)
Cherenkov light effect?

## Slide 17
CTOF Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Plan to remove vertex time dependence with ad hoc functional correction – work in progress

## Slide 18
DC Status
Status and Plans:
Position resolutions 350-450 mm (typical)
Investigate different td function forms to better describe the data
Understand parameter correlations/limits
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Develop status table for reconstruction and MC matching
Alignment (nearly done)
Recalibration monitoring metrics understood

## Slide 19
Forward Tracking Efficiency
Fall-off of tracking efficiency with beam current seen in RG-A, -K, and -B
Working to understand sources of loss; improvements expected moving forward
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
2/16/19
figure of merit
0.5%/nA

## Slide 20
CVT Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Repairs to BMT made in Jan.
SVT and MM performance good for RG-B (N.B. SVT leakage currents)
Alignment of CVT (MM and SVT) in progress – converging soon
Occupancies in SVT < 1% - should yield high efficiency, but this is not the case -
  Possible deficiencies in tracking algorithm ??
Production data @ 35 nA
Zero field alignment data

## Slide 21
SVT Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
500
400
300
200
nA
60
50
40
30
20
10
nA
leakage currents
beam current
Feb. 13 Feb. 18 Feb. 23 Feb. 28
Feb. 13 Feb. 18 Feb. 23 Feb. 28
L1
L2
L6
L4
L3
L5

## Slide 22
Central Tracking Efficiency
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Fall-off of tracking efficiency with beam current
Working to understand sources of loss to see if improvements are possible
2/16/19
figure of merit
1%/nA
positive tracks
negative tracks

## Slide 23
HTCC Calibration Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Gain matching uniformity
Typical Timing Calibration
Average NPhe Response
HTCCy
HTCCx

## Slide 24
LTCC Status
Gain calibrations completed for LTCC
  hit cluster position and energy reconstruction not ready yet
Timing calibration algorithm needs to be developed and tested
Detector Configurations:
Spring run: S2, S3, S5, S6; only S5 was (partially) filled with C4F10
Fall run: S3, S5; only S3 was (partially) filled with C4F10
Gas levels of C4F10 varied with time in unknown way
Winter run: S3, S5 filled with C4F10
Minimal leak rate in both boxes
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 25
RICH Calibration Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
First version of the RICH time calibration suite ready
  total number of parameters is 25024 time offsets and 1564 time walk constants
tested only on run 4013 (software tag 5b.7.4)
DT = meas. – expected
before correction
after correction
Avg. res
<s> = 0.65 ns

## Slide 26
RICH Cherenkov Angle Reconstruction
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Events with 1 electron in the RICH
On-time photons, direct detection (no reflections)
Cherenkov angle [rad]
PMT 9
channel 4416
s(q) ~ 5 mrad
Single photon resolution close to the goal
Expected improvement from misalignment corr.
Reflected photon reconstruction in progress

## Slide 27
CND Status
Calibration suite fully functional
Work in progress:
  Study resolution limitations
  Develop neutral identification in EB
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
dt = 156 ps
Missing neutrons
CND neutrons
= efficiency~8.2%
run 3842
p(e,e'p+)X

## Slide 28
BAND Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Facing Downstream
Side View
beam
Angular coverage: q ~ 155o-176o
Timing resolution < 300 ps
Design neutron efficiency ~35% at 2 MeVee deposited energy
Design neutron momentum resolution 1.5% at 450 MeV/c
BAND Specs:

## Slide 29
BAND Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
  Attenuation corrections
  (L-R) timing offsets
  Speed of light in individual bar
  Paddle-to-paddle offsets from laser
  (Still to come: time-walk and path length corrections)
Reconstructed BAND hits including:
n (250 MeV/c)

## Slide 30
Performance Status
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
System | Spec | Achieved | Spec | Achieved
BAND | <effn> = 35% | TBD | dt < 300 ps | TBD
CND | <effn> = 10% | TBD | dt = 130 ps | 160 ps
CTOF | dt = 65 ps | 130 ps |  | 
DC | dx = 250 – 400 mm | 350 – 500 mm |  | 
ECAL | sE/E = 10%/ | 10%/ | dt < 500 ps | < 600 ps
FT | sE/E < 2%/ + 1% | 3.3%/ | dt < 300 ps | ~200 ps
FTOF | 60 – 100 ps (p1b) | 100 – 150 ps (p1b) | 90 – 160 ps (p1a) | 110 – 250 ps (p1a)
HTCC | effp < 1% | < 1% | <nphe> = 16 | 16
LTCC | eff = 90% | TBD | dt = 1 ns | TBD
RICH | dt < 1 ns | 0.7 ns | p/K rej > 500 | TBD
SVT | S/N > 10 | ~15 | dx = 50 – 65 mm | ~75 mm
System | Spec | Achieved | Spec | Achieved
BAND | <effn> = 35% | TBD | dt < 300 ps | TBD
CND | <effn> = 10% | TBD | dt = 130 ps | 160 ps
CTOF | dt = 65 ps | 130 ps |  | 
DC | dx = 250 – 400 mm | 350 – 500 mm |  | 
ECAL |  |  | dt < 500 ps | < 600 ps
FT |  |  | dt < 300 ps | ~200 ps
FTOF | 60 – 100 ps (p1b) | 100 – 150 ps (p1b) | 90 – 160 ps (p1a) | 110 – 250 ps (p1a)
HTCC | effp < 1% | < 1% | <nphe> = 16 | 16
LTCC | eff = 90% | TBD | dt = 1 ns | TBD
RICH | dt < 1 ns | 0.7 ns | p/K rej > 500 | TBD
SVT | S/N > 10 | ~15 | dx = 50 – 65 mm | ~75 mm
*Entries highlighted in red have not yet met spec's

## Slide 31
Remaining Calibration Work
DC:
  Complete detector alignment
  Change to new t → d function to improve flexibility
FTOF:
  Include vertex correction to RF time
  Implement position-dependent time-walk correction
  Include FADC/TDC hit matching
  Understand systematics in MIP peak position (e.g. vs. torus polarity)
CTOF/CND:
  Include vertex correction to ST (using optimal vertex)
  Correct for time shift associated with event at end of bar
LTCC:
  No timing calibration – need code development
BAND:
  Implement time-walk corrections to timing
  Include path-length corrections
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Short term
Longer term

## Slide 32
Remaining Calibration Work
ECAL:
  Complete studies of long-term energy calibration drift
  Understand systematics in MIP peak position (e.g. vs. torus polarity)
  Develop PMT gain corrections to account for loss with time
  Develop new timing calibration scheme and implement into calibrations
  Understand systematics in timing vs. PID, path length, geom
FT:
  CAL: Refine fitting procedure for energy calibration
  CAL: Include TW correction to optimize timing
  HODO: Address small bias in fits to DE
  TRK: Reconstruction code development underway
CVT:
  Complete alignment work
RICH:
  Complete development of PID scheme and implement in Event Builder
  Complete alignment work
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019
Short term
Longer term

## Slide 33
Subsystem | Group Leader | RG-A/K Calibrator(s)
CND | Silvia Niccolai | Pierre Chatagnon
CTOF | Daniel S. Carman | Chan Kim
DC | Mac Mestayer | Dilini Bulumulla, Latif Kabul, Shirsendu Nanda
ECAL | Cole Smith | Cole Smith/Joshua Artem Tan
FT | Raffaella De Vita | Raffaella De Vita, Nick Zachariou
FTOF | Daniel S. Carman | Christopher McLauchlin
HTCC | Youri Sharabian | Nick Markov, Will Phelps
LTCC | Maurizio Ungaro | Maurizio Ungaro, Burcu Duran
MVT | Maxime Defurne | Maxime Defurne, Guillaume Christiaens
RICH | Marco Contalbrigo | Marco Contalbrigo, Andrey Kim
SVT | Yuri Gotra | Yuri Gotra
Run Group A/K – Calibration Team
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 34
Run Group B – Calibration Team
Subsystem | Group Leader | Calibrator(s)
BAND | Larry Weinstein (ODU) | Florian Hauenstein (ODU), Efrain Segarra (MIT), Reynier Cruz Torres (MIT)
CND | Silvia Niccolai (Orsay) | Paul Naidoo (Glasgow)
CTOF | Daniel S. Carman (JLab) | Achyut Khanal (FIU)
DC | Mac Mestayer (JLab) | Shirsendu Nanda (MSU)
ECAL | Cole Smith (JLab) | Cole Smith, Susan Schadmand (Juelich)
FT | Raffaella De Vita (INFN) | Raffaella De Vita, Nick Zachariou (Edinburgh) Susan Schadmand (Juelich), Alessandra Filippi (INFNTUR)
FTOF | Daniel S. Carman (JLab) | Jose Carvajal (FIU)
HTCC | Youri Sharabian (JLab) | Isabella Illari (GWU)
LTCC | Maurizio Ungaro (JLab) | Maurizio Ungaro
MVT | Maxime Defurne (Saclay) | Maxime Defurne, Gerry Gilfoyle (RICH)
RICH | Marco Contalbrigo (INFN) | Hyon-Suk Jo (KNU)
SVT | Yuri Gotra (JLab) | Yuri Gotra
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019

## Slide 35
Detector Calibration Summary
The calibration suites for all subsystems are well advanced
  Development work continues (optimization and fine-tuning in most cases)
Initial calibrations have been completed for all subsystems
  For most subsystems, the calibrations will need to be redone due to improvements in the reconstruction, alignment, and algorithms
  A complete pass-0 (in progress) studied with the available monitoring tools will be used to determine when calibrations need to be redone
  Understand metrics for recalibration for each subsystem are being developed
Limitations of calibrations toward achieving design specifications being investigated
  Most issues understood and work plans to overcome them in progress
  Working closely with Software Group and RG-A Coordinator to track the remaining work necessary before pass-1 of RG-A spring data
D.S. Carman, CLAS Collaboration Meeting – Mar. 2019