## Slide 1
Sub-ps bunch length measurements
at the JLab FEL
with coherent transition radiation (CTR)
and “Martin-Puplett” interferometer
Pavel Evtushenko, JLab
The key principles of the diagnostic
Operation of the Matrin-Puplett interferometer
An examples of data
Low frequency cut-off
Data Evaluation in frequency domain
Pulsed beam vs. CW beam measurements

## Slide 2
ps and sub-ps bunch length measurements using CTR
Transition radiation is produced when the electron bunch passes a boundary of two media.
Respond time is zero. Shape of the radiation pulse is a “copy” of the electron bunch shape.
When the wave length of the radiation becomes more than the bunch length the radiation becomes COHERENT. (   L )
Power is proportional to:
incoherent radiation  N
coherent radiation  N2
Measurements of the radiation spectrum give information about the bunch length.

## Slide 3
Coherent radiation power spectrum
- Coulomb field of electron bunch
- Fourier transform of the Coulomb field
- Power spectrum of the bunch
- Power spectrum of a single electron
- Bunch Form Factor (BFF)

## Slide 4
The Martin-Puplett interferometer operation

## Slide 5
The Martin-Puplett interferometer

## Slide 6
The Martin-Puplett interferometer
the autocorrelation function is
measured with the help of the MPI
The Wiener-Khintchine theorem says:
“the Fourier transform of the autocorrelation function
is the power spectrum”.
longitudinal field profile
at the MPI entrance
longitudinal field profile
at the MPI exit
detectors measure
intensity IE2

## Slide 7
The MPI scan

## Slide 8
Interferogram - the normalized difference

## Slide 9
The use of two detectors: raw data vs. interferogram

## Slide 10
Bunch length reconstruction
the Gaussian shape
of the bunch is assumed
its power spectrum is
also Gaussian
low frequency cut-off
diffraction on the Golay
cell input window
two filter functions
were considered:
The fit function is used

## Slide 11
Jlab FEL CTR interferogram
step size – 5 um, 128 steps

## Slide 12
Jlab FEL CTR spectrum and the fit function

## Slide 13: JLab FEL (layout and longitudinal matching)
Requirements on phase space:
Long bunch in linac
high peak current (short bunch) at FEL
  bunch length compression at wiggler
“small” energy spread at dump
  energy compress while energy recovering
  “short” RF wavelength/long bunch
   get slope and curvature right
E
f
E
f
E
f
E
f
modified Martin-Puplett
interferometer (step scan)
is used with CTR;
only tune (pulsed) beam
Michelson interferometer
(rapid scan)
is used with CSR;
CW beam

## Slide 14: Interferometers
We use two different interferometers;
essentially both are a modification
of the Michelson interferometer.
The two interferometers differ in
implementation;
Beam splitter
Polarizer
Detector
Focusing element
Mirror position measurements
Modified Marin-Puplett interferometer:
(step scan device)
beam splitter & polarizer (wire grids)
detector (Golay cell)
focusing (Plano-convex lens)
mirror position is set by
step motor
Used with CTR
Michelson interferometer:
(rapid scan device 2 sec/scan)
beam splitter (silicon)
detector (pyroelectric)
focusing (parabolic mirrors)
mirror position is measured by
another built-in interferometer
Used with CSR

## Slide 15: Pulsed beam measurements; RMS 148 fs
The mod. Martin-Puplett interferometer (the Happek device) measurements

## Slide 16: CW beam measurements 0.31 mA; RMS 147 fs

## Slide 17: CW beam measurements 0.62 mA; RMS 151 fs

## Slide 18: CW beam measurements 1.25 mA; 163 fs

## Slide 19: CW beam measurements 2.51 mA; 145 fs

## Slide 20: Pulsed beam measurements vs. CW beam
The bunch length does not change with beam current.